eSupplement II: Methods for collecting and analysing fibrinogen and C-reactive protein

The National Centre for Social Research provides details of the methods for collecting and assaying fibrinogen and C-reactive protein in documentation accompanying results of each of the annual Health Surveys for England.1-3 This supplement summarises these methods and results of quality control.

# Methods for obtaining and analysing blood specimens

After obtaining consent, research nurses collected 4.5ml into a citrate tube for fibrinogen and 6ml blood into a vacutainer with no additive for C-reactive protein estimation and sent them to the laboratory within 24 hours. They were stored at room temperature until analysed.

Measurement of fibrinogen was carried out in 1998 and 2003 by the Department of Haematology at Royal Victoria Infirmary, Newcastle upon Tyne, using the Organon Teknika MDA 180 analyser, using a modification of the Clauss thrombin clotting method. In 1994, fibrinogen was analysed at the West Middlesex University Hospital using a IL ACL2000 analyser, using the nephelometric (clot turbidity) method.

Measurement of high sensitivity C-reactive protein was carried out in 1998 and 2003 by the Biochemistry Department at Royal Victoria Infirmary, Newcastle upon Tyne using the N Latex CRP mono Immunoassay on the Behring Nephelometer II Analyzer.

All analyses were carried out according to Standard Operating Procedures by State Registered Medical Laboratory Scientific Officers under the supervision of the Senior Medical Laboratory Scientific Officers. A schedule of Planned Preventative Maintenance was used for each item of equipment, and was carried out jointly by the manufacturers and the laboratories. Both laboratories undertook internal and external quality control of their analyses of fibrinogen and C-reactive protein.

# Internal quality control

Internal quality control was performed every day either between routine samples or at the end of a batch of routine samples, every 4 hours or 100 specimens. A quality control sample with a known concentration (target value) was assayed many times to establish the variability in the results expressed as coefficient of variation (standard deviation as a percentage of the mean). If the assay precision did not meet acceptability criteria, it was repeated.

In 1994, mean assayed values of fibrinogen and coefficients of variation were reported every for every two months and only for one target value. For 1998 and 2003 mean assayed values and coefficients of variation were reported for each month and two or more types of control materials at different concentrations were used to ensure internal quality over a wider range of concentrations.

Table eII.1 shows results of internal quality control for fibrinogen and Table eII.2 shows results of internal quality control for C-reactive protein. Table eII.3 summarises coefficients of variation for each of the years and overall.

External quality control

The external quality control scheme was operated by the United Kingdom National External Quality Scheme. Samples of known concentration were sent to all participating laboratories, each blind to the concentration of the analyte in the sample. The sample was analysed and the result reported back to the scheme operator. Target values were average of results obtained from all laboratories.

Table eII.4 shows the results of external quality control for fibrinogen and Table eII.5 shows the results of external quality control for C-reactive protein. Table eII.6 summarises the percentage differences between value and target assays for each of the years and overall.

# References

(1) 	National Centre for Social Research, Department of Health. Health Survey for England 1994. 1996. London, The Stationery Office.

(2) 	National Centre for Social Research, Department of Health. Health Survey for England 1998 Cardiovascular Disease. Volume 2: Methodology & Documentation. 1999. The Stationery Office.

(3) 	National Centre for Social Research. Health Survey for England 2003 Volume 3 Methodology and Documentation. Proston K, Primatesta P, editors. 2004. London, The Stationery Office.

eTable II.1. Results of internal quality control of fibrinogen assays, 1994, 1998 and 2003

|  |  | Target value (g/l) | Mean assayed value (g/l) | Coefficient of variation (%) |
| --- | --- | --- | --- | --- |
| 1994 | January | 2.90 | 2.80 | 5.7 |
|  | March | 2.70 | 2.70 | 5.5 |
|  | May | 2.89 | 2.65 | 5.0 |
|  | July | 2.89 | 2.58 | 4.0 |
|  | September | 2.89 | 2.86 | 5.8 |
|  | November | 2.89 | 2.80 | 6.0 |
|  | January 1995 | 3.08 | 3.20 | 3.7 |
| 1998 | March | 1.30 | 1.31 | 9.82 |
|  |  | 2.70 | 2.93 | 8.89 |
|  | April | 1.30 | 1.26 | 8.19 |
|  |  | 2.70 | 2.71 | 6.86 |
|  | May | 1.30 | 1.20 | 8.60 |
|  |  | 2.70 | 2.48 | 6.60 |
|  | June | 1.30 | 1.28 | 8.33 |
|  |  | 2.70 | 2.67 | 6.52 |
|  | July | 1.30 | 1.21 | 8.08 |
|  |  | 2.70 | 2.61 | 6.23 |
|  | August | 1.30 | 1.25 | 8.19 |
|  |  | 2.70 | 2.72 | 6.06 |
|  | September | 1.30 | 1.26 | 11.01 |
|  |  | 2.70 | 2.69 | 7.94 |
|  | October | 1.30 | 1.28 | 9.49 |
|  |  | 2.70 | 2.48 | 6.19 |
|  | November | 1.30 | 1.27 | 10.14 |
|  |  | 2.60 | 2.58 | 6.61 |
|  | December | 1.30 | 1.20 | 10.22 |
|  |  | 2.60 | 2.49 | 7.88 |
|  | January 1999 | 1.30 | 1.33 | 8.16 |
|  |  | 2.60 | 2.66 | 8.06 |
|  | February | 1.30 | 1.28 | 9.40 |
|  |  | 2.60 | 2.63 | 7.52 |
| 2003 | January | 1.50 | 1.49 | 6.7 |
|  |  | 3.00 | 3.07 | 5.5 |
|  | February | 1.50 | 1.45 | 7.3 |
|  |  | 3.00 | 3.09 | 7.5 |
|  | March | 1.30 | 1.39 | 7.6 |
|  |  | 2.90 | 2.88 | 6.5 |
|  | April | 1.30 | 1.42 | 9.1 |
|  |  | 2.90 | 2.97 | 6.0 |
|  | May | 1.45 | 1.46 | 5.3 |
|  |  | 2.90 | 2.98 | 5.1 |
|  | June | 1.45 | 1.46 | 8.2 |
|  |  | 2.90 | 3.03 | 6.4 |
|  | July | 1.45 | 1.43 | 6.1 |
|  |  | 2.90 | 2.97 | 4.3 |
|  | August | 1.45 | 1.49 | 5.6 |
|  |  | 2.90 | 3.03 | 7.5 |
|  | September | 1.45 | 1.43 | 11.7 |
|  |  | 2.90 | 3.03 | 7.0 |
|  | October | 0.90 | 0.99 | 7.7 |
|  |  | 2.90 | 2.89 | 5.1 |
|  | November | 1.45 | 1.04 | 7.6 |
|  |  | 2.90 | 3.10 | 6.2 |
|  | December | 1.45 | 1.03 | 7.9 |
|  |  | 2.90 | 3.00 | 5.3 |
|  | January 2004 | 0.90 | 1.00 | 6.5 |
|  |  | 3.00 | 2.90 | 4.7 |
|  | February | 0.90 | 0.99 | 8.9 |
|  |  | 3.00 | 3.06 | 6.8 |
|  | March | 0.90 | 0.98 | 8.6 |
|  |  | 3.00 | 3.13 | 5.1 |

eTable II.2. Results of internal quality control of C-reactive protein assays, 1998 and 2003

|  |  | Target value (mg/l) | Mean assayed value (mg/l) | Coefficient of variation (%) |
| --- | --- | --- | --- | --- |
| 1998 | February | 1.7 | 1.80 | 4.5 |
|  |  | 16.8 | 16.99 | 4.0 |
|  |  | 40.4 | 41.38 | 3.0 |
|  |  | 91.0 | 95.85 | 4.6 |
|  | March | 1.7 | 1.68 | 2.4 |
|  |  | 16.8 | 16.92 | 3.8 |
|  |  | 40.4 | 40.58 | 2.8 |
|  |  | 91.0 | 92.71 | 5.0 |
|  | April | 1.7 | 1.70 | 3.7 |
|  |  | 16.8 | 16.81 | 3.8 |
|  |  | 40.4 | 40.23 | 3.3 |
|  |  | 91.0 | 90.39 | 3.7 |
|  | May | 1.7 | 1.76 | 4.2 |
|  |  | 15.8 | 15.72 | 4.1 |
|  |  | 27.1 | 27.17 | 3.9 |
|  |  | 37.6 | 37.67 | 3.1 |
|  | June | 1.7 | 1.79 | 5.2 |
|  |  | 15.8 | 15.93 | 3.8 |
|  |  | 27.1 | 27.64 | 3.6 |
|  |  | 37.6 | 38.27 | 2.8 |
|  | July | 1.7 | 1.76 | 5.1 |
|  |  | 15.8 | 15.68 | 2.9 |
|  |  | 27.1 | 27.35 | 2.3 |
|  |  | 37.6 | 38.30 | 2.7 |
|  | August | 1.7 | 1.79 | 5.1 |
|  |  | 15.8 | 15.87 | 2.7 |
|  |  | 27.1 | 27.91 | 2.3 |
|  |  | 37.6 | 39.19 | 2.8 |
|  | September | 1.7 | 1.70 | 4.7 |
|  |  | 15.8 | 15.62 | 3.8 |
|  |  | 27.1 | 27.79 | 2.6 |
|  |  | 37.6 | 39.24 | 2.7 |
|  | October | 1.7 | 1.69 | 3.9 |
|  |  | 15.8 | 15.32 | 2.8 |
|  |  | 27.1 | 27.58 | 2.5 |
|  |  | 37.6 | 38.76 | 3.0 |
|  | November | 1.7 | 1.69 | 4.8 |
|  |  | 15.8 | 15.34 | 3.7 |
|  |  | 27.1 | 27.57 | 3.0 |
|  |  | 37.6 | 38.69 | 2.8 |
|  | December | 1.7 | 1.69 | 4.8 |
|  |  | 15.8 | 15.50 | 4.2 |
|  |  | 27.1 | 27.41 | 3.9 |
|  |  | 37.6 | 38.41 | 3.4 |
|  | January 1999 | 1.5 | 1.48 | 5.6 |
|  |  | 15.8 | 14.95 | 4.8 |
|  |  | 27.1 | 26.61 | 3.8 |
|  |  | 37.6 | 37.47 | 3.5 |
|  | February | 1.5 | 1.50 | 4.4 |

eTable II.2 (continued)

|  |  | Target value (mg/l) | Mean assayed value (mg/l) | Coefficient of variation (%) |
| --- | --- | --- | --- | --- |
| 1998 |  | 15.8 | 15.85 | 3.3 |
|  |  | 27.1 | 27.59 | 3.0 |
|  |  | 37.6 | 38.11 | 3.2 |
| 2003 | January 2003 | 1.40 | 1.52 | 5.3 |
|  |  | 15.00 | 14.78 | 1.5 |
|  |  | 59.80 | 58.32 | 2.1 |
|  | February | 1.40 | 1.44 | 7.6 |
|  |  | 15.00 | 14.21 | 6.8 |
|  |  | 59.80 | 56.41 | 4.8 |
|  | March | 1.40 | 1.39 | 11.3 |
|  |  | 15.00 | 14.15 | 5.0 |
|  |  | 59.80 | 57.16 | 5.9 |
|  | April | 1.40 | 1.43 | 3.5 |
|  |  | 15.00 | 14.28 | 2.2 |
|  |  | 59.80 | 56.29 | 4.1 |
|  | May | 1.40 | 1.45 | 4.1 |
|  |  | 14.80 | 15.28 | 4.5 |
|  |  | 47.20 | 47.83 | 1.1 |
|  | June | 1.40 | 1.40 | 0.0 |
|  |  | 14.80 | 45.36* | 3.5 |
|  |  | 47.20 | 45.36 | 2.4 |
|  | July | 1.40 | 1.48 | 5.4 |
|  |  | 14.18 | 15.12 | 3.2 |
|  |  | 47.20 | 48.81 | 3.5 |
|  | August | 1.40 | 1.54 | 5.9 |
|  |  | 14.80 | 15.40 | 4.7 |
|  |  | 47.20 | 49.75 | 4.1 |
|  | September | 1.50 | 1.40 | 6.4 |
|  |  | 15.00 | 14.93 | 3.2 |
|  |  | 48.00 | 47.53 | 2.7 |
|  | October | 1.50 | 1.38 | 5.1 |
|  |  | 15.00 | 14.85 | 3.2 |
|  |  | 48.00 | 48.23 | 3.8 |
|  | November | 1.50 | 1.50 | 4.7 |
|  |  | 15.00 | 15.13 | 4.1 |
|  |  | 48.00 | 49.56 | 4.3 |
|  | December | 1.50 | 1.45 | 6.9 |
|  |  | 15.00 | 14.71 | 3.6 |
|  |  | 48.00 | 48.71 | 3.1 |
|  | January 2004 | 1.50 | 1.45 | 5.5 |
|  |  | 15.00 | 14.67 | 4.2 |
|  |  | 48.00 | 48.69 | 4.6 |
|  | February | 1.50 | 1.50 | 4.7 |
|  |  | 15.00 | 14.83 | 3.0 |
|  |  | 48.00 | 49.32 | 4.0 |
|  | March | 1.50 | 1.45 | 3.5 |
|  |  | 15.00 | 14.83 | 3.0 |
|  |  | 48.00 | 48.10 | 3.5 |

*Transcribed as reported by the National Centre for Social Research; thought to be typographical error.

eTable II.3: Summary statistics for coefficients of variation for within-assay measurements for fibrinogen and C-reactive protein, 1994, 1998 and 2003

|  | Coefficient of variation (%) |
| --- | --- |
|  |  | Number of measurements | Median | Interquartile range | 90th centile |
| Fibrinogen | 1994 | 7 | 5.5 | 1.8 | 6.0 |
|  | 1998 | 24 | 8.1 | 2.4 | 10.1 |
|  | 2003 | 30 | 6.6 | 2.1 | 8.8 |
|  | All years | 61 | 6.9 | 2.2 | 10.1 |
| C-reactive protein | 1998 | 52 | 3.7 | 1.4 | 4.8 |
|  | 2003 | 45 | 4.1 | 1.8 | 6.4 |
|  | All years | 97 | 3.8 | 1.7 | 6.4 |

###### eTable II.4. Results of external quality control for fibrinogen assays 1994, 1998 and 2003

|  |  | Target value (g/l) | Mean assayed value (g/l) |
| --- | --- | --- | --- |
| 1994 | June | 3.30 | 3.67 |
|  | November | 1.23 | 1.41 |
| 1998 | January | 2.70 | 2.70 |
|  | May | 2.55 | 2.60 |
|  | November | 2.15 | 2.00 |
|  | January 1999 | 2.50 | 2.70 |
| 2003 | January | 1.18 | 1.10 |
|  |  | 2.37 | 2.30 |
|  | February | 2.33 | 2.20 |
|  |  | 2.58 | 2.80 |
|  | March | 0.69 | 0.60 |
|  |  | 1.02 | 0.90 |
|  | April | 1.37 | 1.10 |
|  |  | 1.21 | 1.20 |
|  | May | 2.04 | 1.90 |
|  |  | 1.54 | 1.50 |
|  | June | 1.06 | 1.00 |
|  |  | 2.50 | 2.64 |
|  | July | 2.00 | 2.10 |
|  |  | 1.37 | 1.30 |
|  |  | 1.71 | 1.60 |
|  | August | 1.67 | 1.71 |
|  |  | 2.25 | 2.45 |
|  | September | 1.68 | 1.60 |
|  |  | 1.49 | 1.37 |
|  | October | 1.85 | 1.82 |
|  |  | 1.30 | 1.30 |
|  | November | 2.33 | 2.60 |
|  |  | 1.56 | 1.70 |
|  | December | 2.76 | 2.90 |
|  | January 2004 | 2.07 | 1.96 |
|  |  | 2.87 | 2.94 |
|  |  | 2.37 | 2.30 |
|  | February | 1.97 | 2.03 |
|  |  | 1.49 | 1.49 |
|  | March | 2.07 | 1.98 |

eTable II.5. Results of external quality control for C-reactive protein, 1998 and 2003

|  |  | Target value (mg/l) | Assayed value (mg/l) |
| --- | --- | --- | --- |
| 1998 | January | 37.5 | 45.0 |
|  | February | 62.3 | 79.4 |
|  | March | 72.7 | 88.5 |
|  | April | 82.6 | 87.7 |
|  | May | 97.4 | 108.0 |
|  | June | 117.4 | 118.0 |
|  | August | 147.2 | 152.0 |
|  | September | 11.5 | 13.5 |
|  | October | 21.7 | 24.2 |
|  | November | 33.6 | 36.6 |
|  | December | 59.3 | 59.3 |
|  | January 1999 | 79.6 | 82.0 |
|  | February | 80.2 | 85.4 |
|  | March | 22.6 | 26.4 |
| 2003 | January 2003 | 8.2 | 7.9 |
|  |  | 4.6 | 4.4 |
|  | February | 6.6 | 6.3 |
|  |  | 4.4 | 4.2 |
|  | March | 5.6 | 5.3 |
|  |  | 3.0 | 2.8 |
|  | April | 4.4 | 4.2 |
|  |  | 1.2 | 1.2 |
| 2003 | May | 2.9 | 2.9 |
|  |  | 4.4 | 4.2 |
|  | June | 3.0 | 2.9 |
|  |  | 7.3 | 7.1 |
|  | July | 4.8 | 4.4 |
|  |  | 2.9 | 2.8 |
|  | August | 2.6 | 2.8 |
|  |  | 1.7 | 1.7 |
|  | September | 3.1 | 3.3 |
|  |  | 2.3 | 2.4 |
|  | October | 5.0 | 5.1 |
|  |  | 1.6 | 1.5 |
|  | November | 4.1 | 4.2 |
|  |  | 2.9 | 2.9 |
|  | December | 2.3 | 2.2 |
|  |  | 2.2 | 2.1 |
|  | January 2004 | 5.4 | 5.4 |
|  |  | 5.5 | 5.3 |
|  | February | 1.7 | 1.8 |
|  |  | 1.3 | 1.3 |
|  | March | 6.7 | 6.7 |
|  |  | 3.5 | 3.4 |

eTable II.6: Summary statistics for % differences between value and target assay results for fibrinogen and C-reactive protein, 1994, 1998 and 2003

|  | Difference between value and target results (% of value result) |
| --- | --- |
|  |  | Number of measurements | Median | Interquartile range | 90th centile |
| Fibrinogen | 1994 | 2 | 11.0 | 2.7 | 13.0 |
|  | 1998 | 4 | 1.0 | 8.4 | 7.4 |
|  | 2003 | 30 | -2.9 | 9.0 | 8.0 |
|  | All years | 36 | -1.2 | 10.8 | 8.2 |
| C-reactive protein | 1998 | 14 | 9.0 | 11.7 | 18.0 |
|  | 2003 | 30 | -3.2 | 4.8 | 4.9 |
|  | All years | 44 | 0 | 10.1 | 14.4 |