Supplemental Table 1.  mPCR assessments of methylation are reproducible across two different CE platforms for both normal and full mutation alleles.  Results for four alleles in three samples obtained from eight replicates for both instruments are shown.

| Platform | Sample | Allele CGG # | Average (%) | StDev (%) | CV | 95% CI |
| --- | --- | --- | --- | --- | --- | --- |
| 3130xl | AS10017 | 30 | 74 | 5.3 | 7% | 70.4-77.7 |
|  | AS10017 | 32 | 25 | 2.0 | 8% | 23.4-26.2 |
|  | NA07862 | ~550 | 101 | 10.0 | 10% | 94.5-108.4 |
|  | NA09237 | ~940 | 97 | 15.7 | 16% | 86.5-108.3 |
| 3500xL | AS10017 | 30 | 76 | 5.7 | 8% | 71.8-79.7 |
|  | AS10017 | 32 | 27 | 2.1 | 8% | 25.9-28.8 |
|  | NA07862 | ~550 | 101 | 8.0 | 8% | 95.7-106.8 |
|  | NA09237 | ~940 | 97 | 12.0 | 12% | 88.6-105.3 |