Supplemental Data 3

Details on the MALDI MS sample preparation, data acquisition and pre-processing

Polypeptides were enriched from 5 µL of serum sample using a semi-automated protocol based on reversed phase pre-packed tips (C18 ZipTips).  A CyBi™-Disk robot (CyBio AG, Jena, Germany) equipped with a 96-piston head for 25µL-tips was adapted and used for this purpose.  After C18 ZipTip purification, enriched polypeptides were eluted from the ZipTips, mixed with CHCA matrix and spotted in four replicates of 0.8µL onto a 600µm-AnchorChip™ target plate (Bruker Daltonics) and allowed to dry at room temperature.

Mass spectrometric serum polypeptide profiles were acquired on an Ultraflex II MALDI-TOF/TOF mass spectrometer (Bruker Daltonics, Coventry, UK) in the linear mode.  The overall performance of the mass spectrometer was thoroughly checked every 2-3 weeks using peptide standards and a tryptic digest of bovine serum albumin.  In addition, a commercial human serum standard (Sigma, # S7023) (previously prepared using the ZipTip protocol, aliquotted and frozen) was used prior to each run for a general systems check.  Using this standard serum sample, the resolution and intensity of five major peaks across the mass range of 1,800 to 8,200 Da were checked against previously obtained data and, if necessary, the laser power was adjusted to keep the intensity and resolution of these peaks within their predetermined range of values of ± 2 SD.

All samples were analysed at least in triplicate on separate days.  The same commercial, pooled serum standard as above was freshly prepared 3 times within each batch of serum samples as an internal quality control to assess intra- and inter-run variations.  ClinProTools software V2.1 (Bruker Daltonics) was used for this purpose and allowed to calculate the coefficient of variation (CV) of each of the individual peak areas (more than 100 peaks per spectrum) as well as the mean CV and standard deviation for all of the peaks within the same run or between runs.

Furthermore, the following criteria were checked for each sample:  (1) for each run at least 3 out of the 4 spot replicates contained 1,000 laser shots and were not rejected on grounds of low analyte signal or excessively high matrix signal and (2) at least 2 different runs satisfied condition (1).  However, all samples yielded sufficiently reproducible spectra and none of the samples was excluded.

More information can be found in the SOP below:

Standard Operating Protocol for Blood Serum Polypeptide Profiling Using MALDI-TOF Mass Spectrometry

(January 2009, University of Reading)

Contents

Table of contents………………………………………………………………………………………………………………………….…………….....2

1.	Materials	3

1.1.	Equipment	3

1.2.	Consumables	3

1.2.1.	Blood collection and preparation (in clinical lab)	3

1.2.2.	Sub-sampling (in clinical or MS lab)	3

1.2.3.	Polypeptide extraction, purification and analysis (in MS lab)	3

2.	Methods	5

2.1.	Collecting blood samples and preparing serum	5

2.2.	Sub-sampling of serum samples	6

2.2.1.	Preparations for sub-sampling	6

2.2.2.	Sub-sampling from straws	6

2.2.3.	Sub-sampling from parent microtitre plate	7

2.3.	Preparing solutions	8

2.4.	Extracting polypeptides from serum	8

2.4.1.	Preliminary preparations	9

2.4.2.	Aliquotting ACN	9

2.4.3.	Aliquotting 0.1% TFA	10

2.4.4.	Preparation of plates, prior to extraction	10

2.4.5.	Sample clean-up and MALDI sample preparation	10

2.4.6.	MALDI calibrant and control sample preparation	11

2.5.	MALDI MS data acquisition and analysis	12

2.5.1.	Data acquisition	12

2.5.2.	Data evaluation, processing and preliminary analysis	13

2.5.2.1 Quality control and data filtering/processing	13

2.5.2.2 Preliminary Data Analysis	14

3.	Annexes	15

# Materials

## Equipment

- Scissors

- Pipettors (Fisher Scientific, UK, cat. nos. PMP-124-020E, PMP-124-040V, PMP-124-080J, PMP-124-120A, PMP-133-010G)

- Barcode reader and software (Barcode Solutions, UK, cat. no. Z3071WA)

- Centrifuge

- Freezer at -80(C

- CyBi(-Disk liquid handling robot (CyBio, Northern Europe, cat. no. 0501115-00) housed in a cabinet

- Humidifier unit (Fakir, Germany, Ultronic LB10)

- Mini-centrifuge (Fisher Scientific, UK, cat. no. TEC-100-010Y)

- Vortex Mixer (Fisher Scientific, UK, cat. no. FB15024)

- Ultraflex MALDI-TOF mass spectrometer and FlexControl/FlexAnalysis software (Bruker Daltonics, Germany)

- AnchorChip-600 384-spot MALDI target plates (Bruker Daltonics, Germany, cat. no.209513)

- Laminar Flow Safety Cabinet

- Desk- or lap-top computer

## Consumables

#### Blood collection and preparation (in clinical lab)

  - Vacutainer SST tube (BD, USA, cat. no. 367988)

  - Cryovials (straws)

  - Pipettor tips: 10 µl, 100 µl, 1000 µl (Fisher Scientific, UK, cat. nos. PMP-885-501W, PMP-885-507K, PMP-885-511T)

  - Cryoboxes for sample storage (Fisher Scientific, UK, cat. no. FB71035)

  - Natural Multi PCR tubes: 0.2ml and 1.5ml (Bioquote, UK, cat. nos. 16950F and 16130)

  - Ice

  - Laboratory timer (Fisher Scientific, UK, cat. no. FB51458)

#### Sub-sampling (in clinical or MS lab)

  - Human serum, pooled (Sigma-Aldrich, UK, cat. no. S7023)

  - Pipettor tips: 10 µl, 100 µl, 1000 µl (Fisher Scientific, UK, cat. nos. PMP-885-501W, PMP-885-507K, PMP-885-511T)

  - CyBi( tips (CyBio, Northern Europe, cat. no. OL3800-25-533-N)

  - Cryoboxes for sample storage (Fisher Scientific, UK, cat. no. FB71035)

  - Natural Multi PCR tubes: 0.2ml and 1.5ml (Bioquote, UK, cat. nos. 16950F and 16130)

  - Thermostrips + lids (Abgene Ltd, UK, cat. no. AB-0773)

  - 96-well microtitre plates (Fisher Scientific, UK, cat. no. TUL-962-060A)

  - Universal microtitre plate lids (Camlab Ltd, UK, cat. no. WN\77041001)

  - Ice

  - Virkon( disinfection solution (Fisher Scientific, UK, cat. no. HYG-205-030J)

#### Polypeptide extraction, purification and analysis (in MS lab)

  - Acetonitrile Chromasolv Plus for HPLC (Sigma-Aldrich, UK, cat. no. 34998-1L)

  - Acetone Chromasolv (Sigma-Aldrich, UK, cat. no. 650501-1L)

  - Ethanol 99.8% (Fisher Scientific, UK, cat. no. E/0665/DF/15)

  - (-cyano-4-hydroxycinnamic acid (CHCA) (Bruker Daltonics, Germany, cat. no. 201344)

  - Human serum, pooled (Sigma-Aldrich, UK, cat. no. S7023)

  - Protein and peptide calibration standards (Bruker Daltonics, Germany, cat. nos. 206355 and 222570)

  - Trifluoroacetic acid (TFA) spectrophotometric grade (Sigma-Aldrich, UK, cat. no. 302031-100ml)

  - Water HPLC grade (Rathburn Chemicals, UK, cat. no. RH1020)

  - Pipettor tips: 10 µl, 100 µl, 1000 µl (Fisher Scientific, UK, cat. nos. PMP-885-501W, PMP-885-507K, PMP-885-511T)

  - CyBi( tips (Cybio, Northern Europe, cat. no. OL3800-25-533-N)

  - ZipTips-C18 (Millipore, supplied by Fisher Scientific, UK, cat. no. FDR-597-040G)

  - Ice

  - Modular reservoir, quarter module (Beckman Coulter, UK, cat. no. 372788)

  - Isofreeze PCR rack (Alpha-Laboratories, UK, cat. no. LW5990Y)

  - 384-well plates (Greiner Bio-One, UK, cat.no. 781280X)

  - Teflon reagent reservoirs (CyBio, Northern Europe, cat. no. OL3391-146-24)

  - Small plastic bags and rubber bands for plate storage

  - Natural Multi PCR tubes of 0.2ml and 1.5ml (Bioquote, UK, cat. nos. 16950F and 16130)

  - 96-well microtitre plates (Fisher Scientific, UK, cat. no. TUL-962-060A)

  - Universal microtitre plate lids (Camlab Ltd, UK, cat. no. WN\77041001)

  - pH 0-6 indicator sticks (Fisher Scientific, UK, cat. no. FB33005)

# Methods

Disclaimer:  This protocol has been developed and tested for use only in a research laboratory, not in a clinical environment.  No claims are made for its usefulness, accuracy or safety, and no liability can be accepted for any damages, losses or other expenses of any nature whatsoever arising from its use or supply.

Users of this protocol should have read and understood all relevant Risk and COSHH Assessments and SOPs, in particular health and safety rules and regulations and SOPs for the handling of blood and blood-derived samples such as SOP039 “Safety Considerations in the Handling of Serum Samples from the MRC Project” as used in the MRC-funded project G0301107 at the University of Reading.

## Collecting blood samples and preparing serum

WARNING:  For all blood and blood-derived samples, always observe precautions:  Handle all biological samples as a potential source of pathogens/hazards, use the appropriate protective attire (lab coats, safety glasses, nitrile gloves etc.) and dispose of all (bio)hazardous materials in an appropriate manner.

- Collect venous blood in a Vacutainer tube and label it with the corresponding barcode.

- Gently invert the tube five times to mix the clot activator with the blood.  Allow blood to clot for 1 h at room temperature keeping the tube in the upright position.

NOTE:  Blood should be allowed to clot for exactly 1 h.  If not, artificial differences in the mass spectra of peptides will result from different clotting times of the samples.

- Keep the Vacutainer tubes on wet ice in the upright position until the serum separates (up to 2h).

- Label cryovials (straws) with the corresponding barcode.

- Spin Vacutainer tubes at 1,400 × g for 10 min at room temperature.

- Transfer the serum (upper phase) to the appropriately labelled straws.  The volume of serum per straw should be approximately 500 µl.

- Immediately store all samples at –80 (C.

- Transport samples to the MS laboratory on dry ice.  On arrival they must immediately be stored at –80 (C again.

NOTE:  Avoid freeze-thaw cycles at all steps because they induce proteolysis (in the absence of anti-proteases) and the precipitation of peptides/proteins.

## Sub-sampling of serum samples

NOTE:  Before aliquotting serum samples, they must be randomised to avoid any systematic error due to the sample order or history.

#### Preparations for sub-sampling

WARNING:  All sub-sampling must be carried out in an adequate safety cabinet.  Equipment and consumables used for the purpose of sub-sampling should be housed inside the cabinet and should be labelled with biohazard labels (PCR tubes, thermostrips, solutions, scissors etc.).  Before starting work, ensure that there are two 1-litre Virkon( solutions available that are less than one week old and still pink.  One of the solutions should be placed in a wide-mouthed screw cap jar for disposal of plastic consumables, and the other should be in a screw cap bottle for dealing with accidental spills.

The amount of time that individual samples are outside of the -80°C freezer must be kept to a minimum.  The barcode reader should be plugged in and ready to be used before samples are removed from the freezer.  Lab books and any other stationery should be kept outside the safety cabinet with any pens, markers etc. that might be needed.  The sample label file should be opened on a computer.

Consumables needed for each sample include 1 × 1.5ml-PCR tube, 1 × thermostrip and lid, up to 20 × 200µl-PCR tubes.  A 96-well microtitre plate (parent plate) is used for up to 90 sample aliquots to be run on a liquid handling robot.  The remaining 6 wells are used for control samples as detailed later.  All sample vials, tubes and plates must be kept on wet ice during the entire sub-sampling procedure.

#### Sub-sampling from straws

- Remove the first straw from the freezer, and scan the barcode on the straw directly into the sample label file, and check that the numbers appearing in this file are the same as those written beside the straw bar code.

- Using the scissors, cut the straw just below the closure at the top (green plug end down).  Dispose of the cut-off end of the straw in the Virkon( solution for disposables.

- Invert the tube so that the green plug is now at the top and the other end in a 1.5mL-PCR tube.  Cut the straw again, just below the white section and dispose of the cut-off piece in the Virkon( solution.  Wash the scissors by spraying the blades with ethanol from a wash bottle.

- Lift the straw slightly to allow the thawed serum to drain into the PCR tube.  Use the pipettor with a tip set at 50 µl to blow air through the tube from the top down, in order to expel the last few drops of serum into the tube.  Place the empty straw into the Virkon( solution.  Use the pipettor with the same tip still attached to mix the contents of the PCR tube, to ensure that the whole sample is thoroughly thawed and homogeneous.  Dispose of the pipette tip in the Virkon( solution.

- Transfer 50 µl of serum to the corresponding well in the parent microtitre plate.  From the remaining serum volume sub-sample as many aliquots as are required into 200µl-PCR tubes or into thermostrips (these can be later used for manual or single point analysis).  A volume of 50 µl of sample will later provide sufficient aliquots for 9 child plates (5 µl per plate).

NOTE:  The overall available sample volume determines the number of child plates that can be generated.  Each sample preparation requires a volume of 5 µl of sample, so a total sample volume of 50 µl will realistically provide not more than 9 child plates since 50 µl/5 µl = 10 if the entire volume can be utilised - which is, however, practically impossible.

- Ensure that all of the tube lids are firmly pushed down, then change gloves.

- Record the details of the sample code, date, and number and size of aliquots taken.

- Label all non-plate aliquots with the correct barcode and store them at -80 (C.

- Repeat steps 9-16 for a total of up to 90 straws (i.e. one microtitre plate-load).

- Add 50 µl aliquots of the standard pooled serum sample (thawed and homogenised) to 3 of the 6 empty wells on the microtitre plate.  This standard serum will be used as initial control for the whole run.  The 3 remaining wells are left empty so that there are free spots on the MALDI target plate.  A post-preparation control sample is later spotted onto these empty sample spots (see SST solution, step 31).

NOTE:  Serum must be kept on ice during all aliquotting steps.  Prior to use, the standard serum sample should be aliquotted at 200 and 500 µl into PCR tubes and stored at -80 (C until required.

#### Sub-sampling from parent microtitre plate

NOTE:  The following example of sub-sampling utilises a CyBi™-Disk robot (CyBio, Northern Europe) equipped with a 96-piston head for 25µl-tips and 10 microtitre plate positions and specifically set-up for this task (see Annexe 1).  Thus some steps will most likely differ when using another liquid handling robot.

- Open the customised method “Sub-sampling” from the CyBi( control software.  This method allows replicating the parent plate into 9 child plates transferring 5 µl each time (see Annexe 2A).

- Load a new CyBi( tips rack.

- Place 9 empty 96-well microtitre plates on the CyBi(-disk platform at positions 2 to 10.

- Place the serum parent plate on the platform at position 1 (see Annexe 1).

- Check the simulation of the “Sub-sampling” method for errors, and then run it.

- Cover all the plates with labelled lids and secure the lids with a rubber band.  Put each plate inside a small clear plastic bag, and place the 9 bags together with a copy of the plate plan in a labelled sample box.  Store at -80(C until required.

## Preparing solutions

- 2% TFA:  Mix 98 ml of HPLC grade water with 2 ml of TFA.  Check that the pH is ca. 0.5.

- 0.1% TFA:  Mix 95 ml of HPLC grade water with 5 ml of 2% TFA.  Check that the pH is ca. 2.

- Elution solution (50% ACN/0.1% TFA):  Mix 50 ml ACN with 5 ml TFA 2% and 45 ml HPLC grade water.

- Matrix solution (0.5 mg/ml CHCA in ethanol:acetone; 2:1):  Add 1 ml of ethanol to an 1.5ml-PCR tube containing 3 mg of CHCA and mix until the CHCA is dissolved.  Transfer this solution to a 20ml-Teflon bottle, and add 4 ml ethanol and 1 ml acetone.  Vortex the solution and keep at 4 (C in the dark until needed.

NOTE:  The solutions of 2% TFA, 0.1% TFA and 50% ACN/0.1% TFA should be replaced weekly.  The matrix solution should be freshly prepared on the day of the experiment.

- Virkon( solution:  Add 10 g of Virkon( powder to 1 l of water.  Mix thoroughly until completely dissolved to form a pink coloured solution.  Replace the solution when the pink colour has gone.

- Polypeptide Calibrant Standard (PCS) solution:  Add 125 µl of the 0.1% TFA solution to a vial of peptide calibration standard II and vortex thoroughly to solubilise the peptides.  Do the same with a vial of protein calibration standard I.  Transfer 25 µl from the first vial to a PCR tube and add all the contents of the second vial plus 600 µl of 0.1% TFA.  Mix and aliquot at 4 µl.  Store the aliquots at -20(C (see Annexe 3 for a list of PCS components).

- System Suitability Test (SST) solutions:  A 96-well plate containing 5 µl of the standard serum sample in each well is prepared using the protocol detailed in section 2.4, i.e. steps 32-63, for 96 samples.  The 96 eluates are pooled, mixed and then aliquotted at 4 µl. Aliquots are stored at -80 (C until required.

NOTE:  SST solutions are used as controls to check both the reproducibility of performance of the mass spectrometer (see Section 2.5.1, step 77, and Annexe 6) and the sample preparation (see Section 2.5.2.1).

## Extracting polypeptides from serum

NOTE:  This protocol utilises a CyBi™-Disk robot (CyBio, Northern Europe) equipped with a 96-piston head for 25µl-tips and 10 microtitre plate positions and set-up for this task (see Annexe 1).  Some steps in the protocol will therefore differ when using another liquid handling robot.  Duplicate MALDI target plates are spotted with the prepared samples to provide a back-up in case of problems during the MS acquisition or one-off irregular MALDI sample crystallisations.

#### Preliminary preparations

- Turn on the CyBi™-Disk robot.

- If the ambient relative humidity is less than 50%, increase relative humidity to around 50%, e.g. turn on a humidifier, set it to 50%, turn on a circulating fan in the CyBi( cabinet and close the cabinet door.

- Ensure the solvents and other solutions are less than 1 week old and check that the MALDI targets are clean (follow the manufacturer’s protocols for cleaning the MALDI targets).  Place the targets on the platform at positions 8 and 10, covered with the lid.

- Check that the isofreeze racks are frozen and keep them in the freezer until required.

- Prepare the matrix solution and keep it in the fridge until needed.

- Transfer the ZipTips to their special CyBi( holder.

- Print the 96-well sample plate plan and mark the positions on the CyBi( magazine that are to be kept clear for SST spotting.  Replace the corresponding tips with sawn-off tips so that nothing is aspirated and dispensed by them.

NOTE:  Ensure that all details of the sample preparation are recorded in the lab book, especially noting any deviation from this protocol and important observations made during the procedure.

#### Aliquotting ACN

- Open the method “setup ACN” in the CyBi( Control software (see Annexe 2B).

- Load CyBi( tips for pipetting ACN.

- Place a Greiner 384-well plate on the platform at position 3 (see Annexe 1).

- Place the empty bulk reservoir on the platform at position 9 and carefully fill with ca. 20 ml ACN.

- Check the method “setup ACN” for errors and run it.

- Cover the 384-well plate with a lid until needed.

- Carefully remove the bulk reservoir and dispose of the excess ACN to the waste solvent bottle.

NOTE:  Tips and plates used for ACN aliquotting can be reused for the same purpose.

#### Aliquotting 0.1% TFA

- Open the method “setup TFA” in the CyBi( Control software (see Annexe 2C).

- Load CyBi( tips for pipetting TFA.

- Place a Greiner 384-well plate on the platform at positions 4 and 5 (see Annexe 1).

- Place the empty bulk reservoir on the platform at position 9 and carefully fill with ca. 20 ml of 0.1 % TFA.

- Check the method “setup TFA” for errors and run it.

- Cover both 384-well plates with lids until needed.

- Carefully remove off the bulk reservoir and dispose of the excess 0.1% TFA in the sink.

- Keep the CyBi( tips for the next step.

#### Preparation of plates, prior to extraction

- Pour ca. 5 ml of the 2% TFA solution into the front section of a modular reservoir.  Using a multi-channel pipettor, pipette 10 µl of the 2% TFA solution into each well of a 96-well microtitre plate.  Place the plate on the platform at position 1 and cover it with a lid until needed.

#### Sample clean-up and MALDI sample preparation

- Record the batch name of the sample (plate) to be run (including the date of preparation) on the plate plan and make a copy of the plan.

- Remove the sample plate to be run from the freezer and place it on the platform at position 2.

NOTE:  If for any reason the run is delayed, keep the sample plate on an isofreeze pack, covered with a lid until ready to use.

- Open the method “Sample Run” in the CyBi( Control software and check it for errors.

- Remove the lids from the plates on the CyBi(-Disk platform.

- Run the method “Sample Run” (see Annexe 2D for details).

- When prompted, remove the CyBi( tips, discard them, and load the ZipTips in their special holder.

- After loading the samples, the ZipTips are washed with 0.1% TFA.  During this washing step prepare the eluent plate as follows:  Pour ca. 5 ml elution solution into the front section of a modular reservoir.  From this reservoir use a multichannel pipettor to pipette 7 µl into each well of the 96-well plate and when prompted place the plate on the platform at position 6.

- During the subsequent elution step prepare the matrix plate as follows:  Pour ca. 5 ml matrix solution into the front section of a modular reservoir.  From this reservoir use a multichannel pipettor to transfer 18 µl into each well of the 96-well plate.  Cover the plate with a lid and place it on an isofreeze pack until needed.

- When prompted to change tips, remove and discard the ZipTips.  Replace with a new set of CyBi( tips, where the positions to be used for SST samples have been replaced by sawn-off tips.  Place the matrix plate on the platform at position 7.  Remove the lids from the matrix plate and MALDI targets.

- Once the CyBi(-Disk robot has sampled the eluates, remove the eluate plate from the platform, cover it with a pre-labelled lid and place it on an isofreeze pack.  Mark the lid and plastic bag with “ELUATE”.  Secure the lid with a rubber band, and place the eluate plate back into the bag, together with a copy of the plate plan.  Store the bag in the -80 freezer for future use.

NOTE:  Ensure that the robot cabinet is only opened for the minimum time necessary, to keep the humidity unchanged.  Record the humidity at the time of eluate spotting.

#### MALDI calibrant and control sample preparation

- Keep the robot cabinet closed until the automatically spotted samples are completely dry.

- Take a 4 µl aliquot of each of the SST and PCS solutions from the freezer and allow to thaw briefly.  Spin the solution tubes to ensure that all of the tubes’ contents are at the base of the tubes.

- Transfer 27 µl of the matrix solution from the bottle to two 200µl-PCR tubes, and close the lids tightly.

- Add a 3 µl aliquot of the PCS solution to one of the tubes containing matrix solution, and mix thoroughly.

- On both MALDI target plates spot 0.8 µl of the mixture onto each of 10 random calibrant positions.  These spots are used for external calibration of the MS.  The amount of each polypeptide spotted is ca. 10 fmol for peptides and ca. 50 fmol for proteins.

- Add a 3 µl aliquot of the SST solution to the other tube containing matrix solution, and mix thoroughly.

- On both MALDI target plates spot 0.8 µl of the mixture onto each of the positions (12 per MALDI target plate) marked for SST on the plate plan.  These spots will be used to assess the performance of the MS analysis (see details below in section 2.5.2).

- Close the robot cabinet immediately to maintain the humidity level and keep it closed until the spots are completely dry.

- Turn off the robot, the humidifier and the re-circulating fan inside the robot cabinet (if used).

NOTE:  For thorough calibration across the entire MALDI target plate, the method can be modified to allow automated spotting of PCS solution from a 96-well microtitre plate placed at position 9 on the CyBi(-Disk.

## MALDI MS data acquisition and analysis

#### Data acquisition

NOTE:  This protocol uses an Ultraflex II MALDI-TOF/TOF instrument (Bruker Daltonics) controlled by FlexControl software v.3.0 (Bruker Daltonics) with samples spotted on AnchorChip™ target plates having 384 anchors of 600µm diameter.

- Open the FlexControl software and load optimised FlexControl and AutoXecute methods (see Annexes 4 and 5 for examples as used in the BioCentre at the University of Reading).  These MALDI-TOF settings must be optimised for each individual instrument.

- Take the first target plate from the robot and insert it into the MALDI-TOF mass spectrometer.  The plate should be analysed by the mass spectrometer within 2 h of MALDI sample spotting.  Keep the second target plate covered for possible later use.

- Calibrate the mass spectrometer using a single calibrant spot containing the PCS mixture (see Annexe 3).  Data should be collected automatically by clicking on the tab “Run method on current spot”.

- Collect spectra from an SST spot to check the performance of the mass spectrometer.  The signal-to-noise ratio and resolution of five well-selected peaks should be within their predetermined range of values of ± 2 SD (see Annexe 6).  If this is not the case, re-adjust the laser power accordingly and collect spectra from another SST spot to fulfil these specifications.  Subsequently, data should be collected automatically by clicking on the Tab “Run method on current spot”.

- Load the correct autoXsequence file (containing the names and positions of each sample on the target as well as the file name and path where the spectra will be saved on the hard disk).  This file can be generated by using an Excel spreadsheet from the sample plate plan initially prepared.  Save the file in text format and copy it into the AutoXecute sequence folder on the data drive of the Ultraflex desktop computer.

- Start the automatic run by clicking on the tab “Start automatic run”.  Data collection for a full 384-spot plate takes about 6 h.

#### Data evaluation, processing and preliminary analysis

NOTE:  This protocol uses FlexAnalysis 3.0 and ClinProTools 2.1 software (Bruker Daltonics).

2.5.2.1 Quality control and data filtering/processing

- Open the FlexAnalysis program once the data acquisition is complete.  From the “Process” tab select “Batch check”.  Load the folder containing all acquired spectra and set “the expected number of shots” to 1,000 and click OK.  This will generate a list of any spectra having less than 1,000 shots.  Any spectra with less than 1000 shots should be excluded from further processing.

NOTE:  Poor quality spectra can often result from deterioration of the MALDI target anchors and/or poor crystallisation.  Utilising the second MALDI target plate by acquiring spectra from the spots that are the duplicates of the spots that gave poor quality spectra at step 80 can sometimes give acceptable data for these samples.

- Within the directory containing the MALDI files, create a sub-directory and name it “1000shots”.  Transfer any spectra that gave less than 1000 shots to this directory.  These spectra are excluded from further data processing and analysis.

- Open the ClinProTools program to give a display with four quadrants.  From the File tab, select “Open Model Generation Class” and browse to find the sub-directory containing all spectra files of the SST samples.  Each of the SST spectra is then loaded into the software.  Individual or average spectra are depicted in the top left quadrant, and all of the loaded spectra are depicted in a pseudo-gel view in the bottom left quadrant.

NOTE:  The bottom left quadrant of the ClinProTools display depicts the result of clustering while the top right quadrant gives information about different models (see section 2.5.2.2).

- Select “Open Model Generation Class” again and this time browse to find the sub-directory containing all the spectra given by the standard serum samples prepared with the batch of real samples.  These spectra will be depicted in the pseudo-gel view alongside the spectra of the SST samples, thus allowing a visual comparison of their spectra.  The 2 classes should give virtually identical spectra, in which case the overall preparation of the batch of samples is deemed acceptable.  If the spectra are markedly different then further investigation will be needed to determine the cause.

- Each sample preparation should yield a maximum of 4 spectra, depending on the number of spectra excluded at step 81 (if any).  The spectra for each individual sample are loaded as a class into ClinProTools, and are visually inspected to determine any obvious differences within that class.  If differences are found it may be necessary to exclude further individual spectra on the basis of poor quality.

NOTE:  If the number of replicate spectra per sample falls below 3 as a result of spectra exclusion, it may be necessary to repeat the sample preparation.

2.5.2.2  Preliminary Data Analysis

- ClinProTools can also be used to identify peaks that are significantly different (in intensity) between 2 or more classes of spectra, by using the clustering feature.  It is also possible to use the program to generate other classification models based on lists of discriminatory peaks.  However, this higher level of data analysis and data mining is beyond the scope of this protocol.

- ClinProTools is just one example of the large number of different bioinformatics programs that can be used for the more detailed analysis of the data.  For compatibility with these programs spectra can be exported as tab-separated text files that contain the mass list with the corresponding intensities.  This export facility is available within FlexAnalysis.

# Annexes

Annexe 1.	CyBi(-Disk plates plan

| Position on CyBi(-Disk | Sub-sampling | ACN aliquotting | 0.1% TFA aliquotting | Sample Cleaning and spotting |
| --- | --- | --- | --- | --- |
| 1 | Parent 96-well MTP |  |  | 2% TFA, 96-well MTP |
| 2 | Child 96-well MTP |  |  | Sample, 96-well MTP |
| 3 | Child 96-well MTP | 384-well MTP |  | ACN, 384-well MTP |
| 4 | Child 96-well MTP |  | 384-well MTP | 0.1% TFA, 384-well MTP |
| 5 | Child 96-well MTP |  | 384-well MTP | 0.1% TFA, 384-well MTP |
| 6 | Child 96-well MTP |  |  | Eluate, 96-well MTP |
| 7 | Child 96-well MTP |  |  | Matrix, 96-well MTP |
| 8 | Child 96-well MTP |  |  | MALDI target 1 |
| 9 | Child 96-well MTP | Bulk reservoir | Bulk reservoir | Empty* |
| 10 | Child 96-well MTP |  |  | MALDI target 2 |

* This position can be used for additional solutions such as calibrant solutions for automatic large-scale spotting of calibration samples.

Annexe 2. 	CyBi-Disk TM Methods

(A): “Sub-sampling” method used to sub-sample serum from a parent microtitre plate:

- Aspirate 5 µl aliquots of the serum samples from the parent plate at position 1 and dispense them into the equivalent wells of the child plate at position 2.

- Repeat step 1 eight times each time dispensing the sample aliquots into the next sequential child plate (positions 3 to 10).

(B): “Setup ACN” method used to aliquot ACN into a 384-well microtitre plate:

NOTE:  Each well of a 96-well microtitre plate is equivalent to a set of 4 wells (A, B, C, D) on the 384-well microtitre plate.  The aliquot of ACN is delivered to well A and the other 3 wells are left empty to receive the waste during the ZipTip washing.

- Aspirate 96 aliquots of 25 µl of ACN from the bulk reservoir at position 9 and dispense them into the 384-well microtitre plate (position 3 on the CyBi platform) at well A.

- Repeat step 1 so that the total volume in each well A is 50 µl.

(C): “Setup TFA” method used to aliquot ACN into a 384-well microtitre plate:

NOTE:  Each well of a 96-well microtitre plate is equivalent to a set of 4 wells (A, B, C, D) on the 384-well microtitre plate.  The aliquot of 0.1% TFA is delivered to well A and the other 3 wells are left empty to receive the waste during the ZipTip washing.

- Aspirate 96 aliquots of 25 µl of TFA 0.1% from the bulk reservoir at position 9 and dispense them into the first 384-well microtitre plate (position 4 on the CyBi platform) at well A.

- Repeat step 1 so that the total volume in each well A is 50 µl.

- Aspirate 96 aliquots of 25 µl of TFA 0.1% from the bulk reservoir at position 9 and dispense them into the second 384-well microtitre plate (position 5 on the CyBi platform) at well A.

- Repeat step 3 so that the total volume in each well A is 50 µl.

(D): “Sample run” method used to extract polypeptides from serum samples:

- Add 5 µl 2% TFA to the samples in the 96-well MTP and mix 5 times by pipetting 8 µl up and down.

- Replace manually the CyBi( tips with a rack of 96 ZipTips.

- Condition the ZipTips 3 times with ACN by aspirating 10 µl ACN and dispensing it to the ACN waste wells.

- Equilibrate the ZipTips 3 times with 0.1% TFA by aspirating 10 µl of 0.1% TFA and dispensing it to the TFA waste wells.

- Load the samples onto ZipTips by aspirating and dispensing 10 × 10 µl of the sample mixture.  (Aspirating and dispensing of samples needs to be slow because of the viscosity of serum and the back pressure due to the packed phase in the ZipTips).  An overstroke of 1 µl is used at the last dispensing step to expel any remaining drops of serum from the ZipTips.

- Wash the ZipTips by aspirating and dispensing 3 × 10 µl of 0.1% TFA.  Each time the solution is dispensed into a new waste well to avoid cross-contamination.  An overstroke of 1 µl is also used at the last dispensing step to expel any remaining drops of wash solution from the ZipTips.

- Elute the analytes from the ZipTips by aspirating and dispensing the eluent (7 µl of 50% ACN / 01.% TFA) 5 times into the same well of the eluate plate.

- Replace manually the ZipTips with a rack of 96-CyBi( Tips.

- Transfer 2 µl of the eluate to the matrix plate wells.  Mix the eluate with the matrix solution by aspirating and dispensing 5 × 15 µl.

- Aspirate 10 µl of the eluate/matrix mixture solution and spot 0.8 µl onto each of 4 subsequent positions of the MALDI target plate, so that there are 4 replicates for each prepared sample.

Annexe 3.	Polypeptide Calibrant List

| Polypeptide | Ion type | Average mass m/z |
| --- | --- | --- |
| Peptide calibration standard II mixture (Bruker Daltonics, #222570) |  |  |
| Bradykinin Fragment 1-7 | [M+H]+ | 757.858 |
| Angiotensin II | [M+H]+ | 1047.189 |
| Angiotensin I | [M+H]+ | 1297.486 |
| Substance P | [M+H]+ | 1348.642 |
| Bombesin | [M+H]+ | 1620.860 |
| Renin Substrate Tetradecapeptide Porcine | [M+H]+ | 1760.026 |
| ACTH clip 1-17 | [M+H]+ | 2094.427 |
| ACTH clip 18-39 | [M+H]+ | 2466.681 |
| Somatostatin 28 | [M+H]+ | 3149.574 |
| Protein calibration standard I mixture (Bruker Daltonics, #206355) |  |  |
| Insulin | [M+H]+ | 5734.56 |
| Ubiquitin I | [M+2H]2+ | 4283.45 |
| Cytochrome C | [M+2H]2+ | 6181.05 |
| Ubiquitin I | [M+H]+ | 8565.89 |

Annexe 4. 	FlexControl settings optimised for the analysis of polypeptides prepared from serum samples

Annexe 5.

FlexControl autoXecute method optimised for the automatic analysis of polypeptides prepared from serum samples

Annexe 6.

Peak list and parameters used for assessing the performance of the MS analysis (SST) based on a reference spectrum of a standard serum sample

| m/z | Signal/Noise | Resolution |
| --- | --- | --- |
|  | Mean (M) | M-2σ (min) | M+2σ (max) | Mean (M) | M-2σ (min) | M+2σ (max) |
| 1808.82 | 1919 | 1214 | 2624 | 184 | 177 | 190 |
| 2981.54 | 2841 | 2399 | 3283 | 234 | 226 | 241 |
| 5003.29 | 4159 | 3330 | 4988 | 293 | 286 | 300 |
| 6447.50 | 959 | 749 | 1169 | 261 | 224 | 298 |
| 8126.35 | 1344 | 1073 | 1616 | 146 | 114 | 177 |