Supplementary content

The repeatability of blood gas parameters, PCO2 gap, and PCO2 gap to arterial-to-venous oxygen content difference in critically ill adult patients

Jihad MALLAT1 M.D., Ali LAZKANI1 M.D., Malcolm LEMYZE1 M.D., Florent PEPY1 M.D., Mehdi Meddour1 M.D., Stéphanie BARRAILLER1 M.D., Gaëlle GASAN1 M.D., Johanna TEMIME1 M.D., Nicolas VANGRUNDERBEECK1 M.D., Laurent TRONCHON1 M.D., and Didier THEVENIN1 M.D.

Materials and methods

Procedure

GEM® Premier™ 3000:

The GEM Premier 3000 is a portable analyzer used for analysis of whole blood samples at the point-of-care in a clinical setting. This analyzer simultaneously measures blood gas, electrolytes, glucose, lactate, and hematocrit values. Determination of analyte concentration is based on potentiometry (pH, PCO2, Na+, K+ and Ca++), amperometry (PO2, glucose and lactate), and conductometry (hematocrit). Sample volume is 150 µL of whole blood for the full menu of analytes and results are available approximately 90 seconds from sample introduction. The device comprises two components: the instrument and the replaceable cartridge as a primary component. The disposable cartridge is a closed system that contains all components needed for sample analysis: reagents, electrochemical sensor, sampling stylus and the waste container. Cartridge insertion in the analyzer is followed by automated sensor calibration.  Subsequent calibrations are performed automatically at regular scheduled intervals in order to ensure constant accuracy of the analyzer performance. The device performs one-point calibrations on an average of every 20 min and two-point calibrations every 2 h. one-point calibration is also performed automatically after every sample analysis.  Furthermore, GEM Premier 3000 has an active automated performance quality control program termed Intelligent Quality Management, which continually controls the entire process of sample analysis, enables instantaneous error detection, and performs corrective actions for error elimination. Moreover, the analyzer detects insufficient samples, bubbles or microthrombi, and in these cases it will cancel the test and display a message.

The within-run imprecision for aqueous quality control materials [standard deviation (SD) or % coefficient of variation (%CV)] was pH: SD= 0.0020 to 0.0067; PCO2: 1.3% to 2.8%; PO2: 1% to 2.3%; and hematocrit: 0% to 0% [1]. The between-run imprecision (day-to-day precision) for aqueous quality control materials was pH: SD= 0.0064 to 0.0085; PCO2: 2.3% to 2.7%; PO2: 2.0% to 2.3%; and hematocrit: 0% to 0% [1]. Accuracy of the analyzer was good with clinically acceptable correlation in comparison with central laboratory instruments [1-3].

References:

- Bénéteau-Burnat B, Bocque MC, Lorin A, Martin C, Vaubourdolle M. Evaluation of the blood gas analyzer Gem PREMIER 3000. Clin Chem Lab Med 2004; 42: 96-101.

- Steinfelder-Visscher J, Weerwind PW, Teerenstra S, Brouwer MH. Reliability of point-of-care hematocrit, blood gas, electrolyte, lactate and glucose measurement during cardiopulmonary bypass. Perfusion 2006; 21:33-37.

- Jacobs E, Ancy JJ, Smith M. Multi-site performance evaluation of pH, blood gas, electrolyte, glucose, and lactate determinations with the GEM Premier 3000 critical care analyzer. Point of Care 2002; 3: 135-144.