Supplementary Table 1: Common wearable devices on the market.

| Locations | Type of device | Company/Institution | Product name | Parameters | FDA cleared |
| --- | --- | --- | --- | --- | --- |
| Head |  |  |  |  |  |
|  | Headbands | Earable | Earable | EEG, EOG, EMG and sleep | Yes |
|  |  | Zeit | Headband | EEG | No or not sure |
|  |  | Bose | SoundSport Pulse | HR | No or not sure |
|  |  | MIT | Earphones | HR and PPG | No or not sure |
|  | Glasses | Korea University | E-glasses | EEG and EOG | No or not sure |
|  |  | Microsoft | Glabella | Pulse transit time and cuff-less BP | No or not sure |
|  |  | Google | Google Glass | HR and respiratory rate | No or not sure |
| Trunk |  |  |  |  |  |
|  | Patches | iRhythm | iRhythm Zio | HR and ECG | Yes |
|  |  | BardyDX | BardyDX CAMTM | HR and ECG | Yes |
|  |  | VivaLNK | VivaLNK ECG | HR and ECG | Yes |
|  |  | BioTelemetry | BioTel Heart | HR and ECG | Yes |
|  |  | MediBioSense | Vital Patch | HR, PA, temperature, and ECG | Yes |
|  |  | Global Instrumentation | Global Instrumentation M5 | HR and ECG | Yes |
|  |  | Shenzhen Institute of Advanced Technology | 3-in 1 | HR, ECG, and PPG | Yes |
|  | Chest straps | Polar | Polar H10 | HR | Yes |
|  |  | Garmin | Garmin HRM | HR | No or not sure |
|  |  | Wahoo | Fitness TICKR | HR | No or not sure |
|  | T-shirts | Hexoskin | Hexoskin T-shirt | HR, respiratory rate, PA, sleep, and ECG | No or not sure |
|  |  | Emglare | Emglare T-shirt | HR and ECG | No or not sure |
|  |  | AMSU | AMSU T-shirt | HR and ECG | No or not sure |
|  |  | Lenovo | SmartVest | HR and ECG | No or not sure |
| Limbs |  |  |  |  |  |
|  | Watches | Apple | Apple Watch | HR, PA, sleep, and ECG | Yes |
|  |  | Samsung | Galaxy Watch | HR and ECG | Yes |
|  |  | Fitbit | Sense | HR, PA, and sleep | Yes |
|  |  | Biobeat | BB-613WP | HR and cuffless BP | Yes |
|  |  | Omron | HeartGuide | HR, PA, sleep, and cuff-based BP | Yes |
|  |  | Huawei | Watch GT 2 | HR, PA, and ECG | Yes |
|  |  | Withings | ScanWatch | HR, PPG, ECG, and sleep | Yes |
|  |  | Amazfit | Amazfit GTR 3 | HR, SpO2, respiratory rate, and ECG | No or not sure |
|  |  | OPPO | OPPO Watch ECG | HR, PA, ECG, and sleep | Yes |
|  | Bands | AliveCor | KardiaBand | HR and ECG | Yes |
|  |  | Samsung | Galaxy Fit 2 | HR, PA, and sleep | No or not sure |
|  |  | Huawei | Band | HR, SpO2, and PA | No or not sure |
|  |  | Xiaomi | Band | HR, SpO2, PA, and sleep | No or not sure |
|  |  | Fitbit | Charge 5 | HR, SpO2, ECG, PA, and sleep | Yes |
|  |  | Nike | FuelBand | PA | No or not sure |
|  |  | Codoon | Wristband | PA and sleep | No or not sure |
|  |  | Amazon | Halo Wristband | PA, sleep, and body mass index | No or not sure |
|  | Rings | Jakcom | R3 | HR, PA, and sport | No or not sure |
|  |  | Motiv | Ring | HR, PA, and sleep | No or not sure |
|  |  | Oura | Ring | HR, PA, and sleep | No or not sure |
|  | Shoes | iFeel | iFeel iit | PA | No or not sure |
|  |  | NURVV Run | Smart Insoles | PA | No or not sure |

BP: Blood pressure; EEG: Electroencephalogram; EOG: Electrooculogram; EMG: Electromyography; ECG: Electrocardiography; FDA: Food and Drug Administration; HR: Heart rate; PPG: Photoplethysmography; PA: Physical activity; SpO2: Blood oxygen.

Supplementary Table 2: Summary of the validity of various wearable devices compared to criterion measure for PA, HR, and sleep.

| Parameter | Reference | Subjects | Activity | Criterion | Devices | Validity |
| --- | --- | --- | --- | --- | --- | --- |
| PA | Case et al[27] | N = 14 (4 males); 28.1 ± 6.2 years | Walk 500 and 1500 steps twice each at 3.0 mph on a treadmill. | Manual counting | Digi-Walker SW-200, Fitbit Zip, Fitbit One, Fitbit Flex, Jawbone UP24, Nike Fuelband, iPhone 5s Fitbit App, iPhone 5s Health Mate App, iPhone 5s Moves App, Galaxy S4 Moves App. | No specific values were reported, but it can be seen from the figures in the article that for all step trials, the Fitbit One had the best accuracy, followed by the Fitbit Zip and the Digi-Walker SW-200. |
|  | Fokkema et al[28] | N = 31 (15 males); 32 ± 12 years | Participants walked three walking speeds for 10 min each: slow (3.2 km/h), average (4.8 km/h), and vigorous (6.4 km/h) | Manual counting | Polar Loop, Garmin Vivosmart, Fitbit Charge HR, Apple Watch Sport, Pebble Smartwatch, Samsung Gear S, Misfit Flash, Jawbone Up Move, Flyfit, Moves | The most accurate device for the slow walking speed is Fitbit Charge (MAPE = −0.7%), followed by Garmin Vivosmart (MAPE = 1.0%). The most accurate device for average walking speed is Apple Watch (MAPE = 0.0%), followed by Garmin Vivosmart (MAPE = 0.2%). The most accurate device for vigorous walking is Apple Watch (MAPE = 0.5%), followed by Samsung Gear (MAPE = 1.1%). |
|  | Xie et al[29] | N = 44 (22 males): 22.2 ± 2.2 years | Resting, walking, running, cycling, and sleeping | Manual counting | Apple Watch 2, Samsung Gear S3, Jawbone Up3, Fitbit Surge, Huawei Talk Band B3, Xiaomi Mi Band 2, Dongdong App, Ledongli App. | Dongdong App (MAPE: 1 ± 3%) < Ledongli App (MAPE: 2 ± 3%) < Huawei Talk Band B3 (MAPE: 2 ± 4%) < Fitbit Surge (MAPE: 6 ± 9%) < Jawbone Up3 (MAPE: 6 ± 16%) < Xiaomi Mi Band (MAPE: 6 ± 17%) < Samsung Gear (MAPE: 21 ± 33%) < Apple Watch (MAPE: 42 ± 37%). |
|  | Höchsmann et al[30] | N = 20 (6 males); 18–70 years | Participants walked for 5 min at each of the predefined walking speeds of 1.6 km/h, 3.2 km/h, 4.8 km/h, and 6.0 km/h on a treadmill | Manual counting | iPhone SE, Samsung Galaxy S6 Edge, Garmin Vivofit 2, ActiGraph wGTX+. | For treadmill walking ≥ 3.2 km/h, the MAPE of the iPhone SE (all positions) and the Garmin Vivofit was small (<3). For treadmill walking at 4.8 km/h and 6.0 km/h, Samsung Galaxy (MAPE<3%) and ActiGraph (MAPE<3%) showed small MAPE (<3%). The ActiGraph showed high MAPE (17–47) for all walking conditions. Overall, the iPhone SE and the Garmin Vivofit are accurate tools for step counting in different age groups and during various walking conditions. |
|  | Boolani et al[31] | N = 120 (74 males); 21.4 ± 3.7 years | Two 6-minute walks, one at a comfortable pace and the other at a fast pace | Manual counting | Fitbit Zip, Garmin Vivofit, Basis B1 band, Misfit Shine, Runtastic Pedometer, Nike FuelbandS | Fitbit Zip (MAPE = 0.18%, ICC = 0.76, ME = −1.2) < Garmin Vivofit (MAPE = 1.3%, ICC = 0.70, ME = 8.5) < Basis B1 band (MAPE = 2.2%, ICC = 0.60, ME = 14.2) < Misfit Shine (MAPE = 4.7%, ICC = 0.68, ME = 30.6) < Nike FuelbandS (MAPE = 8.6%, ICC = 0.22, ME = 56.4) < Runtastic Pedometer (MAPE = 12.8%, ICC = 0.64, ME = 83.6) |
|  | Navalta et al[32] | N = 20 (12 males); 22.2 ± 5.8 years | Hiking and trail running | Manual counting | Fitbit Surge 2, Garmin Vivosmart, Leaf Health Tracker, Polar A360, Samsung Gear 2, Spire Activity Tracker, Stryd Power Meter. | For 5-min hiking: Garmin Vivosmart (mean bias: −4.38 ± 27.93), Polar A360 (mean bias: 14.18 ± 49.62), Fitbit Surge (mean bias: 18.41 ± 57.08), Leaf health tracker (mean bias: −19.70 ± 30.14), Samsung Gear (mean bias: 20.31 ± 49.69), Stryd Power Meter (mean bias: 22.31 ± 28.89), Spire Activity Tracker (mean bias: 144.81 ± 136.82). For 5-min trail runs: Fitbit Surge (mean bias: 5.00 ± 140.66), Garmin Vivosmart (mean bias: −7.59 ± 20.87), Leaf Health Tracker (mean bias: −14.78 ± 13.86), Stryd Power Meter (mean bias: 16.09 ± 20.53), Samsung Gear (mean bias: 42.53 ± 140.46), Polar A360 (mean bias: 52.36 ± 84.59), Spire Activity Tracker (mean bias: 103.71 ± 242.49) |
|  | Bunn et al[33] | N = 20 (10 males); | According to standards of the Consumer Technology Association (CTA) | Manual counting | Apple iWatch series 1, Fitbit Surge, Garmin 235, Moto 360, Polar A360, Suunto Spartan Sport, Suunto Spartan Trainer, and TomTom Spark 3 | During walking, Moto 360 (MAPE = 1.07%) < Garmin (MAPE = 2.34%) Apple iWatch (MAPE = 3.72%) < Polar A360 (MAPE = 4.60%) < Suunto Trainer (MAPE = 4.99%) < TomTom (MAPE = 10.11%) < FitBit (MAPE = 11.20%) < Suunto Sport (MAPE = 22.93%). During running, Moto 360 (MAPE = 0.32%) < Garmin (MAPE = 0.59%) < Suunto trainer (MAPE = 1.14%) < Suunto sport (MAPE = 2.04%) < TomTom (MAPE = 2.98%) < Apple iWatch (MAPE = 3.67%) < FibBit (MAPE = 5.16%) < Polar A360 (MAPE = 10.66%). |
| Heart rate | Gillinov et al[37] | N = 50 (23 males); 38 ± 12 years | Treadmill, stationary bicycle, elliptical trainer | 12-lead ECG | Polar H7, Scosche Rhythm+, Apple Watch, Fitbit Blaze, Garmin Forerunner 235, TomTom Spark | Polar H7 (rc = 0.996) < Apple Watch (rc = 0.92) < TomTom Spark (rc = 0.83) < Garmin Forerunner (rc = 0.81) < Scosche Rhythm (rc = 0.75) < Fitbit Blaze (rc = 0.67). |
|  | Koshy et al[38] | N = 102 (67 males); 68.0 ± 15.0 years | - | 12-lead ECG | Apple Watch Series 1 and Fitbit Blaze | Bias was the calculated ME between the smart watches-HR and ECG-HR. In sinus rhythm, both devices with a low bias (Fitbit Blaze & Apple Watch, Bias = 1 beat). In atrial arrhythmias, Apple Watch (Bias = −5 beats) < Fitbit Blaze (Bias = −18 beats). |
|  | Boudreaux et al[39] | N = 50 (22 males); 22.71 ± 2.99 years | Graded cycling and three sets of four resistance exercises. | 6-lead ECG | Apple Watch Series 2, Fitbit Blaze, Fitbit Charge 2, Polar H7, Polar A360, Garmin Vivosmart HR, TomTom Touch, and Bose SoundSport Pulse headphones | During graded exercise cycling: Apple Watch (MAPE = 4.14%) < Polar H7 (MAPE = 6.87%) < Bose (MAPE = 7.44%) < TomTom (MAPE = 12.33%) < Polar A360 (MAPE = 19.48%) < Fitbit Blaze (MAPE = 21.6%) < Fitbit Charge (MAPE = 21.36%) < Garmin Vivosmart (MAPE = 25.38%). During resistance exercise: Bose (MAPE = 6.24%) < Polar H7 (MAPE = 6.31%) < Polar A360 (MAPE = 8.66%) < Fitbit Charge (MAPE = 9.79%) < Garmin Vivosmart (MAPE = 10.66%) < Apple Watch (MAPE = 10.99%) < Fitbit Blaze (MAPE = 13.74%) < TomTom (MAPE = 19.14%) |
|  | Thiebaud et al[40] | N = 20 (18 males); 22.0 ± 3 years. | Participants exercised on a treadmill at 3.2 km/h, 4.8 km/h, 6.4 km/h, 8 km/h, and 9.7 km/h for 3 min at each speed with no rest between speeds. | 12-lead ECG | Fitbit Surges, Microsoft Bands, TomTom Cardios | Fitbit Surge (MAPE: 2.17–8.06%), TomTom Cardio (MAPE: 1.01–7.49%), Microsoft Band (MPAE: 1.31–7.37%). |
|  | Xie et al[29] | N = 44 (22 males): 22.2 ± 2.2 years | Resting, walking, running, cycling, and sleeping | Manual counting | Samsung Gear S3, Apple Watch Series 2, Fitbit Surge, Xiaomi Mi Band 2 | Samsung Gear (MAPE: 4 ± 3%) < Apple Watch (MAPE: 7 ± 8%) < Fitbit Surge (MAPE: 8 ± 12%) < Xiaomi Mi Band (MAPE: 12 ± 13%) |
|  | Hwang et al[41] | N = 51 (27 males); 44.1 ± 16.6 years | - | 12-lead ECG | Apple Watch Series 2, Samsung Galaxy Gear S3, and Fitbit Charge 2 | The accuracy of the baseline HR measurement, i.e., within ±5 beats/min of the ECG value: Apple (100%), Galaxy (100%), and Fitbit (100%). The accuracy of the supraventricular tachyarrhythmia HR measurements to within ±5 beats/min of the ECG value: Apple (89.3%), Galaxy (89.7%), and Fitbit (83.3%); |
|  | Thomson et al[42] | N = 30 (15males); 23.5 ± 3.0 years | Three-minute stages with speed and incline increasing every 3 min until volitional fatigue | 12-lead ECG | Fitbit Charge HR 2 and Apple Watch | Relative error rates (RER) were used. Apple Watch (RER: 2.4–5.1%) < Fitbit (RER: 3.9–13.5%) |
|  | Navalta et al[43] | N = 21 (11 males); 31 ± 11 years | Out-and-back trail runs with the first 1.61 km is uphill, and then return 1.61 km is downhill | 1-lead ECG | Garmin Fenix 5, Jabra Elite Sport earbuds, Motiv ring, Scosche Rhythm+, Suunto Spartan Sport watch | Validity was determined through three methods: MAPE, Bland-Altman Limits of Agreement (LOA), and Lin’s Concordance Coefficient (rc). Suunto Spartan Sport (MAPE = 2%, LOA = −62 to 61, rc = 0.96), Scosche Rhythm+ (MAPE = 6%, LOA = −114 to 120, rc = 0.79), Garmin Fenix (MAPE = 13%, LOA = −32 to 162, rc = 0.32), Motiv ring (MAPE = 16%, LOA = −52 to 96, rc = 0.29), Jabra Elite Sport (MAPE = 23%, LOA = −464 to 503, rc = 0.38). |
|  | Baek et al[44] | N = 15 (males); 23.7 ± 3.0 years | Conventional and Nordic walking | 12-lead ECG | Polar H7 and Fitbit Charge 2 | In conventional walking, Polar H7 (AbsD: 1.30 ± 2.12 beats/min) < Fitbit Charge (AbsD: 3.68 ± 4.51 beats/min). In Nordic walking, Polar H7 (AbsD: 1.37 ± 4.87 beats/min) < Fitbit Charge (AbsD: 6.60 ± 6.81 beats/min). Overall, Polar H7 (rc = 0.96) > Fitbit Charge (rc = 0.84). |
| Sleep | De Zambotti et al[55] | N = 65 (37 males); 15.8 ± 2.5 years | - | PSG | Jawbone UP | Mean (or bias) and SD of the differences between Jawbone UP and PSG outcomes were provided. The MEs ± SD of the Jawbone UP for measuring TST, sleep efficiency, total wake time, SOL, and WASO were −10.0 ± 20.5, −1.9 ± 4.2, 9.3 ± 20.4, −1.3 ± 10.9, and 10.6 ± 14.7 min, respectively. |
|  | De Zambotti et al[56] | N = 44 (18 males); age: 19–61 years | - | PSG | Fitbit Charge 2™ | Mean (or bias) and SD of the differences between Fitbit Charge and PSG outcomes were provided. For participants without PLMS (N = 35): TST (−9 ± 24 min), SOL (4 ± 9 min), WASO (5 ± 19 min), time in N1+N2 (light sleep) (−34 ± 34 min), time in N3 (deep sleep) (24 ± 28 min), time in REM (1 ± 27 min). For participants with PLMS (N = 9): TST (−8 ± 39 min), SOL (7 ± 10 min), WASO (2 ± 26 min), time in N1+N2 (light sleep) (−35 ± 49 min), time in N3 (deep sleep) (28 ± 35 min), time in REM (0 ± 37 min). |
|  | De Zambotti rt al. 2019.[57] | N = 41 (28 males); 17.2 ± 2.4 years | - | PSG | ŌURA ring | Mean (or bias) and SD of the differences between ŌURA ring and PSG outcomes were reported. TST (−1.3 ± 21.7 min), SOL (−0.2 ± 7.0 min), WASO (1.5 ± 20.7 min), time in N1+N2 (−3.7 ± 66.2 min), time in N3 (19.6 ± 41.2 min), time in REM (−17.2 ± 50.2 min) |
|  | Chinoy et al[58] | N = 34 (14 males); 28.1 ± 3.9 years. | - | PSG | Fatigue Science Readiband, Fitbit Alta HR, Garmin Fenix 5S, Garmin Vivosmart 3 | Fatigue Science Readiband (bias: 13.3 min for TST, 2.8 min for SE, −0.7 min for SOL, −12.5 min for WASO, NANs for light sleep, deep sleep, and REM); Fitbit Alta HR (bias: 2.6 min for TST, 0.9 min for SE, −3.1 min for SOL, −2.1 min for WASO, 20.0 min for light sleep, −6.0 min for deep sleep, −11.4 min for REM); Garmin Fenix (bias: 43.7 min for TST, 10.6 min for SE, 0.8 min for SOL, −49.5 min for WASO, 29.0 min for light sleep, 6.1 min for deep sleep, −8.6 min for REM) Garmin Vivosmart (bias: 46.8 min for TST, 10.1 min for SE, −1.1 min for SOL, −47.6 min for WASO, 34.7 min for light sleep, 3.7 min for deep sleep, 8.4 min for REM). |

AbsD: Absolute difference; ECG: Electrocardiography; HR: Heart rate; ICC: Intra-class correlation coefficient; MAPE: Mean absolute percent error; ME: Mean difference; PA: Physical activity; PLMS: Periodic limb movement of sleep; PSG: Polysomnography; rc: Lin’s concordance correlation coefficient; REM: Rapid-eye-movement; SOL: Sleep onset latency; TST: Total sleep time; WASO: Wake after sleep onset.

Supplementary Table 3: Summary of the use of wearable devices in cardiovascular care.

| Cardiovascular care | Reference | Year | Wearable Device | Technology | Summary |
| --- | --- | --- | --- | --- | --- |
| Arrhythmia detection | Hannun et al[61] | 2019 | Zio patch (iRhythm Technologies, USA) | Deep neural network | A DNN was trained on 91,232 single-lead ECGs from 53,549 patients to classify 12 rhythm classes by using the Zio monitor (iRhythm Technologies, USA). The DNN achieved an average area under the ROC of 0.97 on a test dataset consisting of 328 ECG records from 328 patients. |
|  | mSToPS Trial[62] | 2018 | Zio patch (iRhythm Technologies, USA) | Cox proportional model | The mSToPS Trial included 3476 individuals in the non-monitored cohort and 2659 in the monitored cohort (1366 in the immediate monitoring group using Zio patch and 1293 in the delayed monitoring group). After 4 months, the incidence of new AF cases was 3.9% (53/1366) in the immediate monitoring group vs. 0.9% (12/1293) in the delayed monitoring group (AbsD, 3.0% [95% CI: 1.8–4.1%]). After 12 months, 1738 of 2659 in the monitored cohort finished follow-up; 190 new cases of AF were detected, 109 of 1738 (6.7 per 100 person-years) in the actively monitored cohort and 81 of 3476 (2.6 per 100 person-years) among observational controls (AbsD, 4.1 [95% CI: 3.9–4.2]). |
|  | Apple heart study[63] | 2019 | Apple Watch | Deep neural network | The study included 419,297 participants and used the Apple Watch PPG sensor to measure changes in blood flow to distinguish irregular pulse; 2161 participants (0.52%) received notifications of an irregular pulse. Among 2161 participants, 450 finished the validation via ECG patches after receiving irregular pulse notifications. AF was identified in 34% participants (97.5% CI: 29 to 39) (139/450). Of the 86 participants who received irregular pulse notifications on Apple devices during simultaneous use of an ECG patch, 72 showed AF on concurrent ECG patch strips. This resulted in a PPV for the irregular pulse notification of 0.84 (95% CI: 0.76 to 0.92). |
|  | Health eHeart Study[64] | 2018 | Apple Watch | Deep neural network | 9750 participants (347 participants with AF) were enrolled in the study. The DNN exhibited a C statistic of 0.97 (95% CI: 0.94–1.00; P < 0.001) to detect AF against the reference standard 12-lead ECG–diagnosed AF in the external validation cohort of 51 patients undergoing cardioversion. In an exploratory analysis relying on self-report persistent AF in ambulatory participants, the C statistic was 0.72 (95% CI: 0.64–0.78). |
|  | Bumgarner et al[65] | 2018 | Apple Watch combination with the AliveCor KardiaBand | - | A total of 100 patients were enrolled (age 68 ± 11 years) in the study. There were 169 simultaneous ECG and Kardia Band recordings; 57 recordings were noninterpretable by the Kardia Band. Compared with ECG, the Kardia Band interpreted AF with 93% sensitivity, 84% specificity, and a K coefficient of 0.77. Among 113 cases where Kardia Band and physician readings of the same recording were interpretable, the agreement was excellent (K coefficient = 0.88). |
|  | Huawei Heart Study[66] | 2019 | Huawei Watch GT, Honor Watch, and Honor Band | PPG algorithm | The study aimed to determine the feasibility of AF screening in a large population-based cohort using smart device-based PPG technology, combined with a clinical care AF management pathway. Among those with PPG monitoring (mean age 35 years, 86.7% male), 424 (mean age 54 years, 87.0% male) received a “suspected AF” notification (424/187,912, 0.23%). Of those effectively followed up, 227 individuals (227/262, 87.0%) were confirmed as having AF, with the PPV of PPG signals being 91.6% (95% CI: 91.5–91.8%). |
|  | Chen et al[67] | 2020 | Amazfit Health Band 1S (Huami Technology, Anhui, China) | Deep learning (SEResNet) | The study assessed PPG and single-channel ECG data collected from 401 patients (251 normal individuals and 150 ECG-diagnosed AF patients) using the Amazfit Health Band 1S. The algorithm was developed using a convolution neural network (SEResNet) based on the RealBeats Artificial Intelligence Biological Data Engine (Huami Technology). The sensitivity, specificity, and accuracy of wristband PPG readings were 88.00%, 96.41%, and 93.27%, respectively, and those of wristband ECG readings were 87.33%, 99.20%, and 94.76%, respectively. When physicians judged the original wristband ECG records, the sensitivity, specificity, and accuracy were 96.67%, 98.01%, and 97.51%, respectively. |
| BP measurement | Watanabe et al[70] | 2017 | Cuffless BP estimation | PPG algorithm | The CLB only uses PPG to estimate BP. Various feature parameters of the pulse wave in 887 participants at rest and under exercise and mental stress were collected and extracted from one PPG sensor. These parameters were then applied to the BP estimation algorithm. A validation test was conducted via simultaneously monitoring BP with the CLB and with a cuff-type auscultatory device. The MAD between the BP value of CLB and the cuff-wearing sphygmomanometer was <8 mm Hg (6.1 for SBP), suggesting that the BP values measured by using CLB meet the Institute of Electrical and Electronics Engineers 1708–2014 standard. |
|  | Moon et al[71] | 2020 | InBodyWATCH | Neural network | 35 adults (age 57.1 ± 17.9 years) were enrolled in the study. The accuracy of BP measurement using InBodyWATCH with an individualized estimation based on a neural network model and using a manual sphygmomanometer was compared. The ME was 2.2 mmHg ± 6.1 mmHg for SBP and −0.2 mmHg ± 4.2 mmHg for DBP; these were not significant (P = 0.472 for SBP and P = 0.880 for DBP). The estimated SBP/DBP ratios obtained from the InBodyWATCH within ±5 mmHg of manual SBP/DBP were 71.4%/83.8%; within ±10 mmHg, they were 86.7%/98.1%; and within ±15 mmHg, they were 97.1%/99.0%. |
|  | Chandrasekhar et al[72] | 2018 | A smartphone | Stepwise regression | A smartphone-based BP monitoring via the oscillometric finger-pressing method was reported. The training data was collected from 31 individuals (age, 31 ± 7 years; height, 170 ± 8 cm; weight, 68 ± 10 kg; 39% females) via a standard automatic arm cuff. Stepwise regression was used to determine the model parameters. The testing data set was collected from 30 new users and five additional experienced users (age, 33 ± 8 years; height, 173 ± 4 cm; weight, 72 ± 5 kg; 0% females). The smartphone-based device yielded bias and precision errors of 3.3 mmHg and 8.8 mmHg for SBP and −5.6 mmHg and 7.7 mmHg for DBP over a 40- to 50-mmHg range of BP. |
|  | Van Helmond et al[73] | 2019 | Everlast watch, BodiMetrics | Statistics | The study analyzed 85 individuals, including 36 (42%) women, 32 (38%) self-reported hypertension, and 97% of these (n = 31) reported taking antihypertensive medications. The Everlast smartwatch and the BodiMetrics performance monitor presented low accuracy and were not competitive as BP measurement devices. The average differences between the Everlast smartwatch and hospital-grade automated sphygmomanometer were systolic BP of 16.9 ± 13.5 mm Hg and diastolic BP of 8.3 ± 6.1 mm Hg. The average difference between the BodiMetrics performance monitor and hospital-grade automated sphygmomanometer was systolic BP of 5.3 ± 4.7 mm Hg. |
| Other diagnostic applications |
| Diabetes, high cholesterol, high BP, and sleep apnea detection | Ballinger et al[74] | 2018 | Fitbit, Apple Watch, and Wear OS | Deep learning (Semi-supervised, multi-task LSTM) | The study trained and validated a semi-supervised, multi-task LSTM on 57,675 person-weeks of data from popular wearable devices such as Fitbit, Apple Watch, or Android Wear, showing high accuracy at detecting multiple medical conditions, including diabetes (0.8451), high cholesterol (0.7441), high BP (0.8086), and sleep apnea (0.8298). |
| Diabetes detection | Avram et al[75] | 2020 | Azumio Smartphone app | Deep neural network | The study developed a DNN to detect diabetes on 53,870 individuals as a primary cohort based on PPG collected from smartphones and validated in 7806 contemporary cohorts and 181 clinic cohorts. The network achieved an area under the curve for prevalent diabetes of 0.766 in the primary cohort (95% CI: 0.750–0.782; sensitivity 75%, specificity 65%) and 0.740 in the contemporary cohort (95% CI: 0.723–0.758; sensitivity 81%, specificity 54%). |
| Hyperkalaemia diagnosis | Galloway et al[76] | 2018 | AliveCor ECG | Deep neural network | Study data included 709,000 patients, 2.1 million ECGs, and 4.0 million serum potassium values. Each patient used an investigational version of a portable, AliveCor smartphone ECG device to acquire a 4-hour ECG recording during two separate dialysis sessions, with concurrent blood testing. The DNN was trained to detect hyperkalemia using only ECG leads I and II. The sensitivity by duration was 94%, at a specificity of 74%. |
| Myocardial infarction | Sopic et al[77] | 2018 | SmartCardia INYU | Machine learning (Random Forest classifications) | Based on random forest classifications, this paper demonstrated a real-time early detection and prevention of myocardial infarction technique on SmartCardia INYU device and evaluated its energy consumption and battery life performance. The participators consisted of 104 individuals (83 in the training set and 21 subjects in the test set). The classifier that uses all available features (n = 72) reaches a geometric mean of 83.26% (Sensitivity = 87.95%, Specificity = 78.82%). |
| Arterial stiffness measurement | Miao et al[78] | 2019 | SIAT 3-in-1 | Machine learning | This study, including valid data from 501 individuals, proposed a wearable sensor for arterial stiffness monitoring via machine learning techniques. The features were extracted from ECG and PPG signals and were selected based on a genetic algorithm. Using Omron arterial stiffness equipment as the reference, the proposed model achieved the best accuracy of 0.89, 0.2136, and 6.2432 in correlation coefficient, ME, and standard difference for vascular age estimation. |
| Cardiovascular risk assessment | Women’s Health Study[79] | 2018 | ActiGraph GT3X+ | Statistics (proportional hazards regression) | The study analyzed 16741 women wearing triaxial accelerometers (ActiGraph GT3X+) for ≥10 h/day on ≥4 days. Via proportional hazards regression, the associations of the total volume of PA (PA, total accelerometer counts per day), MVPA, minutes per day, LPA, minutes per day, and sedentary behavior (minutes per day) with mortality were examined. A strong inverse association between overall volume of PA and all-cause mortality was observed. The magnitude of risk reduction (≈60%–70%, comparing extreme quartiles) was far larger than that estimated from metanalyses of studies using self-reported PA (≈20%–30%). |
|  | Akbulut et al[80] | 2018 | CVDiMo | Machine learning (Multiclass decision forest) | The study proposed a wearable device called CVDiMo which provides continuous medical monitoring and creates a health profile with the risk of disease over time. Six different biosignals (ECG, Body Temperature, Pulse Oximeter, GSR, BP, Glucometer) from two different test groups with 30 participants were analyzed. Besides the patients’ biosignals, using the PA results and stress levels deduced from the emotional state analysis reached a higher performance in risk estimation. The highest accuracy of classifying the short-term health status was 96% on the multiclass decision forest model. |

AbsD: Absolute difference; AF: Atrial fibrillation; BP: Blood pressure; CVDiMo: Cardiovascular Disease Monitoring; CI: Confidence interval; CLB: Cuff-less BP estimation; DBP: Diastolic blood pressure; DNN: Deep neural network; ECG: Electrocardiogram; GSR: Galvanic Skin Response; LPA: Light physical activity; LSTM: Long-short term machine; MAD: Mean absolute difference; ME: Mean difference; MVPA: Moderate to vigorous physical activity; PA: Physical activity; P: P-value; PPG: Photoplethysmography; PPV: Positive predictive value; ROC: Receiver operating characteristic curve; SBP: Systolic blood pressure.