Table S1. CTQ Assay Performance Compared to PIS by Study Site

| Sensitivity | Specificity |
| --- | --- |
| Study Site [clinic type] (prevalence) | Endocervical swabs | Vaginal swabs | Female urine neat | Female urine UPT | Male urethral swabs | Male urine neat | Male urine UPT | Endocervical swabs | Vaginal swabs | Female urine neat | Female urine UPT | Male urethral swabs | Male urine neat | Male urine UPT |
| 1 [STD] (Females: 18.1%) (Males: 23.8%) | 89.5% (17/19) (66.9, 98.7) | 94.7% (18/19) (74.0, 99.9) | 100.0% (19/19) (82.4, 100.0) | 100.0% (19/19) (82.4, 100.0) | 95.8% (23/24) (78.9, 99.9) | 95.8% (23/24) (78.9, 99.9) | 95.8% (23/24) (78.9, 99.9) | 100.0% (86/86) (95.8, 100.0) | 100.0% (86/86) (95.8, 100.0) | 97.6% (83/85) (91.8, 99.7) | 100.0% (85/85) (95.8, 100.0) | 100.0% (77/77) (95.3, 100.0) | 98.7% (76/77) (93.0, 100.0) | 98.7% (76/77) (93.0, 100.0) |
| 2 [STD, FP] (Females: 10.0%) (Males: 0.0%) | 85.7% (6/7) (42.1, 99.6) | 100.0% (7/7) (59.0, 100.0) | 100.0% (7/7) (59.0, 100.0) | 85.7% (6/7) (42.1, 99.6) | NA | NA | NA | 98.4% (62/63) (91.5, 100.0) | 98.4% (62/63) (91.5, 100.0) | 100.0% (63/63) (94.3, 100.0) | 98.4% (62/63) (91.5, 100.0) | 100.0% (21/21) (83.9, 100.0) | 100.0% (21/21) (83.9, 100.0) | 100.0% (21/21) (83.9, 100.0) |
| 3* [STD] (Females: 16.1%) (Males: 24.1%) | 96.0% (24/25) (79.6, 99.9) | 100.0% (25/25) (86.3, 100.0) | 92.0% (23/25) (74.0, 99.0) | 96.0% (24/25) (79.6, 99.9) | 87.8% (43/49) (75.2, 95.4) | 98.0% (48/49) (89.1, 99.9) | 98.0% (48/49) (89.1, 99.9) | 96.2% (125/130) (91.3, 98.7) | 97.7% (127/130) (93.4, 99.5) | 97.7% (127/130) (93.4, 99.5) | 97.7% (127/130) (93.4, 99.5) | 98.1% (151/154) (94.4, 99.6) | 99.4% (153/154) (96.4, 100.0) | 98.1% (151/154) (94.4, 99.6) |
| 4** [FP] (Females: 8.5%) | 90.3% (28/31) (74.2, 98.0) | 96.8% (30/31) (83.3, 99.9) | 93.5% (29/31) (78.6, 99.2) | 93.5% (29/31) (78.6, 99.2) | NA | NA | NA | 98.5% (329/334) (96.5, 99.5) | 100.0% (334/334) (98.9, 100.0) | 100.0% (335/335) (98.9, 100.0) | 99.4% (333/335) (97.9, 99.9) | NA | NA | NA |
| 5*[STD] (Females: 11.0%) (Males: 22.4%) | 88.2% (15/17) (63.6, 98.5) | 100.0% (17/17) (80.5, 100.0) | 82.4% (14/17) (56.6, 96.2) | 82.4% (14/17) (56.6, 96.2) | 100.0% (17/17) (80.5, 100.0) | 100.0% (17/17) (80.5, 100.0) | 100.0% (17/17) (80.5, 100.0) | 98.6% (136/138) (94.9, 99.8) | 99.3% (137/138) (96.0, 100.0) | 100.0% (138/138) (97.4, 100.0) | 99.3% (137/138) (96.0, 100.0) | 98.3% (58/59) (90.9, 100.0) | 100.0% (59/59) (93.9, 100.0) | 96.6% (57/59) (88.3, 99.6) |
| 6** [STD] (Females: 10.0%) (Males: 15.5%) | 100.0% (7/7) (59.0, 100.0) | 100.0% (7/7) (59.0, 100.0) | 100.0% (7/7) (59.0, 100.0) | 100.0% (7/7) (59.0, 100.0) | 90.9% (10/11) (58.7, 99.8) | 100.0% (11/11) (71.5, 100.0) | 100.0% (11/11) (71.5, 100.0) | 96.8% (61/63) (89.0, 99.6) | 96.8% (61/63) (89.0, 99.6) | 100.0% (63/63) (94.3, 100.0) | 100.0% (63/63) (94.3, 100.0) | 96.7% (58/60) (88.5, 99.6) | 98.3% (59/60) (91.1, 100.0) | 98.3% (59/60) (91.1, 100.0) |
| 7 [OB/GYN, FP] (Females: 12.3%) | 88.9% (8/9) (51.8, 99.7) | 77.8% (7/9) (40.0, 97.2) | 88.9% (8/9) (51.8, 99.7) | 88.9% (8/9) (51.8, 99.7) | NA | NA | NA | 100.0% (64/64) (94.4, 100.0) | 100.0% (64/64) (94.4, 100.0) | 100.0% (64/64) (94.4, 100.0) | 100.0% (64/64) (94.4, 100.0) | NA | NA | NA |
| * Testing sites ** Collection sites only An additional site performed Viper testing but did not enroll subjects |

Table S2. Female Results for All Assays

|  | NAAT 1 | NAAT 2 | Viper |  |
| --- | --- | --- | --- | --- |
| Category | PT Endocervical | PT Urine | AC2 Endocervical | AC2 UTT | CTQ Endocervical | CTQ Vaginal | CTQ Neat | CTQ UPT | Symptomatic Status |
|  |  |  |  |  |  |  |  |  | N | Y | Total |
| Positive by FDA required PIS | - | + | - | + | - | - | + | + | 0 | 2 | 2* |
|  | - | + | - | + | - | + | - | - | 1 | 0 | 1* |
|  | - | + | - | + | - | + | + | + | 2 | 4 | 6* |
|  | - | + | - | + | + | - | + | + | 1 | 0 | 1 |
|  | - | + | + | + | - | + | + | + | 1 | 0 | 1 |
|  | - | + | + | + | + | + | + | + | 1 | 3 | 4 |
|  | + | NA | + | + | + | + | + | + | 0 | 1 | 1 |
|  | + | - | + | - | + | + | - | - | 1 | 2 | 3 |
|  | + | - | + | + | + | - | - | - | 0 | 1 | 1 |
|  | + | - | + | + | + | + | - | - | 0 | 1 | 1 |
|  | + | - | + | + | + | + | - | + | 1 | 0 | 1 |
|  | + | - | + | + | + | + | + | - | 1 | 0 | 1 |
|  | + | - | + | + | + | + | + | + | 0 | 2 | 2 |
|  | + | + | - | + | + | + | + | + | 1 | 0 | 1 |
|  | + | + | + | + | + | + | - | + | 1 | 0 | 1 |
|  | + | + | + | + | + | + | + | - | 0 | 1 | 1 |
|  | + | + | + | + | + | + | + | + | 46 | 41 | 87 |
| TOTAL Infected | 115 |
| FDA PIS negative (with at least two positive results with at least one by reference assay) | + | + | - | - | + | + | + | + | 0 | 1 | 1 |
|  | - | - | + | + | + | + | + | + | 0 | 1 | 1 |
|  | - | - | + | + | + | + | - | + | 0 | 1 | 1 |
|  | - | - | + | + | - | - | + | - | 0 | 1 | 1 |
|  | - | - | + | - | + | + | - | - | 1 | 0 | 1 |
|  | - | - | + | - | - | + | - | + | 1 | 0 | 1 |
|  | - | - | - | + | + | + | - | - | 0 | 1 | 1 |
|  | - | - | + | - | + | - | - | - | 1 | 1 | 2 |
| FDA PIS negative TOTAL (Open to interpretation) | 9** |
| FDA PIS negative (1 or less positive result) | NA | - | - | + | - | - | - | - | 1 | 0 | 1 |
|  | - | - | - | - | - | - | - | + | 1 | 2 | 3 |
|  | - | - | - | - | - | - | + | - | 0 | 2 | 2 |
|  | - | - | - | - | - | + | - | - | 0 | 1 | 1 |
|  | - | - | - | - | + | - | - | - | 6 | 2 | 8 |
|  | - | - | - | + | - | - | - | - | 0 | 4 | 4 |
|  | - | - | + | - | - | - | - | - | 1 | 3 | 4 |
|  | - | + | - | - | - | - | - | - | 0 | 1 | 1 |
|  | + | - | - | - | - | - | - | - | 0 | 2 | 2 |
|  | NA | - | - | - | - | - | - | - | 11 | 2 | 13 |
|  | - | NA | - | - | - | - | NA | NA | 0 | 1 | 1 |
|  | - | NA | - | - | - | - | - | - | 1 | 0 | 1 |
|  | - | I | - | - | - | - | - | - | 5 | 1 | 6 |
|  | - | - | NA | - | - | - | - | - | 1 | 2 | 3 |
|  | - | - | - | - | LE | - | - | - | 0 | 1 | 1 |
|  | - | - | - | - | - | NA | - | - | 1 | 0 | 1 |
|  | - | - | - | - | - | - | - | - | 362 | 456 | 818 |
| Total FDA PIS Negative |  |  | 879 |
| *: infections in the urethra only: PT, AC2 and CTQ endocervical swabs are negative **: analysis was done with and without these nine infections since these were open to interpretation |  |  |  |

Table S3. Male Results for All Assays

|  | NAAT 1 | NAAT 2 | Viper |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| PIS CT | PT Urethral | PT Urine | AC2 Urethral | AC2 UTT | CTQ Urethral | CTQ Neat | CTQ UPT | Symptomatic Status |
|  |  |  |  |  |  |  |  | N | Y | Total |
| Positive by FDA required PIS | NA | + | + | + | + | + | + | 0 | 6 | 6 |
|  | - | + | - | + | - | + | + | 3 | 2 | 5 |
|  | - | + | + | - | - | - | - | 0 | 1 | 1 |
|  | - | + | + | + | + | + | + | 1 | 4 | 5 |
|  | + | - | + | + | - | + | + | 0 | 1 | 1 |
|  | + | - | + | + | + | - | - | 0 | 1 | 1 |
|  | + | + | NA | + | + | + | + | 2 | 5 | 7 |
|  | + | + | - | + | + | + | + | 0 | 1 | 1 |
|  | + | + | + | - | + | + | + | 0 | 1 | 1 |
|  | + | + | + | + | - | + | + | 1 | 0 | 1 |
|  | + | + | + | + | + | + | + | 28 | 44 | 72 |
| Total Infected | 35 | 66 | 101 |
| FDA PIS negative (at least two positive results with at least one by reference assay) | + | + | NA | - | + | + | + | 1 | 0 | 1 |
|  | - | - | + | + | - | + | + | 0 | 1 | 1 |
|  | - | - | + | - | + | - | - | 0 | 1 | 1 |
|  | - | - | + | - | + | + | + | 1 | 0 | 1 |
| FDA PIS negative TOTAL (Open to interpretation) | 2 | 2 | 4 |
| FDA PIS negative (1 or less positive result) | NA | - | - | - | - | - | + | 0 | 1 | 1 |
|  | - | - | - | - | - | - | + | 0 | 2 | 2 |
|  | - | - | - | - | + | - | - | 0 | 2 | 2 |
|  | - | - | - | - | + | - | + | 0 | 1 | 1 |
|  | - | - | - | + | - | - | - | 0 | 3 | 3 |
|  | - | - | + | - | - | - | - | 1 | 1 | 2 |
|  | + | - | - | - | - | - | - | 1 | 0 | 1 |
|  | NA | - | - | - | - | - | - | 4 | 12 | 16 |
|  | - | I | NA | - | - | - | - | 1 | 0 | 1 |
|  | - | I | - | - | - | - | - | 4 | 1 | 5 |
|  | - | - | NA | - | - | - | - | 10 | 20 | 30 |
|  | - | - | - | - | - | - | - | 157 | 146 | 303 |
| Total FDA PIS Negative | 180 | 191 | 371 |

Table S4.  CTQ Assay Performance by Urine Volume

| Specimen Type | Urine Volume | N | Sensitivity | Specificity |
| --- | --- | --- | --- | --- |
| Neat Urine | 20 mL - 30.49 mL | 489 | 96.6% (56/58) | (88.1% - 99.6%) | 99.1% (427/431) | (97.6% - 99.7%) |
|  | 30.5 mL - 50.49 mL | 879 | 94.3% (133/141) | (89.1% - 97.5%) | 99.6% (735/738) | (98.8% - 99.9%) |
|  | more than 50.5 mL | 97 | 100.0% (17/17) | (80.5% - 100.0%) | 98.8% (79/80) | (93.2% - 100.0%) |
| UPT Urine | 20 mL - 30.49 mL | 489 | 94.8% (55/58) | (85.6% - 98.9%) | 98.4% (424/431) | (96.7% - 99.3%) |
|  | 30.5 mL - 50.49 mL | 879 | 95.0% (134/141) | (90.0% - 98.0%) | 99.3% (733/738) | (98.4% - 99.8%) |
|  | more than 50.5 mL | 97 | 100.0% (17/17) | (80.5% - 100.0%) | 97.5% (78/80) | (91.3% - 99.7%) |